## SCCP cal_1
SCCP
[13 empty] | 220915_Blank_1 | 220915_1 | 220915_2 | 220915_3 | 220915_4 | 220915_5 | 220915_6 | 220915_7 | 220915_8 | 220915_9
Sample ID |  | STD1 | STD2 | STD3 | STD4 | STD5 | STD6 | STD7 | STD8 | STD9 |  |  | Sample01 | Sample02 | Sample03 | Sample04 | Sample05 | Sample06 | Sample07 | Sample08 | Sample09 | Sample10 | STD5
Sample name | IDL | SCCP-MIX-CS1 | SCCP-MIX-CS2 | SCCP-MIX-CS3 | SCCP-MIX-CS4 | SCCP-MIX-CS5 | SCCP-MIX-CS6 | SCCP-MIX-CS7 | SCCP-MIX-CS8 | SCCP-MIX-CS9 |  | MDL | 220915_9 STD5: SCCP-MIX-CS5
Sample | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 1
Extract-1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 1
Extract-2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 1
LC vial | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 1 | bar: Sample01 to 220915_9 STD5
D18-HBCD (μg/mL) | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 |  | 0.15 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.15
Unit | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL |  | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/mL
C10_Cl4 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl5 | 0.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl6 | 0.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl7 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl4 | 0.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl5 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl4 | 0.01 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl5 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl4 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl4-8 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl4-9 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl4-10 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl4-10 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SCCP |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
%Cl |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | - | bar: Sample01 to 220915_9 STD5
Concentrations of target compounds in sample
Unit | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL |  | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/mL
C10_Cl4 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl5 | 0.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl6 | 0.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl7 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl4 | 0.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl5 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl4 | 0.01 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl5 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl4 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Concentrations of target compounds in final solvent
Unit | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL |  | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL
C10_Cl4 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl5 | 0.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl6 | 0.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl7 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl4 | 0.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl5 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl4 | 0.01 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.01 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl5 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl4 | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.005 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Measurement result
C10_Cl4 | #DIV/0!
C10_Cl5 | #DIV/0!
C10_Cl6 | #DIV/0!
C10_Cl7 | #DIV/0!
C10_Cl8 | #DIV/0!
C11_Cl4 | #DIV/0!
C11_Cl5 | #DIV/0!
C11_Cl6 | #DIV/0!
C11_Cl7 | #DIV/0!
C11_Cl8 | #DIV/0!
C11_Cl9 | #DIV/0!
C12_Cl4 | #DIV/0!
C12_Cl5 | #DIV/0!
C12_Cl6 | #DIV/0!
C12_Cl7 | #DIV/0!
C12_Cl8 | #DIV/0!
C12_Cl9 | #DIV/0!
C12_Cl10 | #DIV/0!
C13_Cl4 | #DIV/0!
C13_Cl5 | #DIV/0!
C13_Cl6 | #DIV/0!
C13_Cl7 | #DIV/0!
C13_Cl8 | #DIV/0!
C13_Cl9 | #DIV/0!
C13_Cl10 | #DIV/0!
D18-HBCD
Calibration solution concentration
Unit | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | Factor
C10_04Cl | 0.05 | 0.00022 | 0.00055 | 0.0011 | 0.0022 | 0.0055 | 0.011 | 0.022 | 0.055 | 0.11 | 0.011
C10_05Cl | 0.02 | 0.00034 | 0.00085 | 0.0017 | 0.0034 | 0.0085 | 0.017 | 0.034 | 0.085 | 0.17 | 0.017
C10_06Cl | 0.02 | 0.00038 | 0.00095 | 0.0019 | 0.0038 | 0.0095 | 0.019 | 0.038 | 0.095 | 0.19 | 0.019
C10_07Cl | 0.005 | 0.0003 | 0.00075 | 0.0015 | 0.003 | 0.0075 | 0.015 | 0.03 | 0.075 | 0.15 | 0.015
C10_08Cl | 0.002 | 0.0001 | 0.00025 | 0.0005 | 0.001 | 0.0025 | 0.005 | 0.01 | 0.025 | 0.05 | 0.005
C11_04Cl | 0.02 | 0.0011 | 0.00275 | 0.0055 | 0.011 | 0.0275 | 0.055 | 0.11 | 0.275 | 0.55 | 0.055
C11_05Cl | 0.005 | 0.00164 | 0.0041 | 0.0082 | 0.0164 | 0.041 | 0.082 | 0.164 | 0.41 | 0.82 | 0.082
C11_06Cl | 0.002 | 0.00136 | 0.0034 | 0.0068 | 0.0136 | 0.034 | 0.068 | 0.136 | 0.34 | 0.68 | 0.068
C11_07Cl | 0.002 | 0.00144 | 0.0036 | 0.0072 | 0.0144 | 0.036 | 0.072 | 0.144 | 0.36 | 0.72 | 0.072
C11_08Cl | 0.002 | 0.001 | 0.0025 | 0.005 | 0.01 | 0.025 | 0.05 | 0.1 | 0.25 | 0.5 | 0.05
C11_09Cl | 0.002 | 0.0003 | 0.00075 | 0.0015 | 0.003 | 0.0075 | 0.015 | 0.03 | 0.075 | 0.15 | 0.015
C12_04Cl | 0.01 | 0.00078 | 0.00195 | 0.0039 | 0.0078 | 0.0195 | 0.039 | 0.078 | 0.195 | 0.39 | 0.039
C12_05Cl | 0.005 | 0.00142 | 0.00355 | 0.0071 | 0.0142 | 0.0355 | 0.071 | 0.142 | 0.355 | 0.71 | 0.071
C12_06Cl | 0.002 | 0.00138 | 0.00345 | 0.0069 | 0.0138 | 0.0345 | 0.069 | 0.138 | 0.345 | 0.69 | 0.069
C12_07Cl | 0.002 | 0.00112 | 0.0028 | 0.0056 | 0.0112 | 0.028 | 0.056 | 0.112 | 0.28 | 0.56 | 0.056
C12_08Cl | 0.002 | 0.00124 | 0.0031 | 0.0062 | 0.0124 | 0.031 | 0.062 | 0.124 | 0.31 | 0.62 | 0.062
C12_09Cl | 0.002 | 0.00072 | 0.0018 | 0.0036 | 0.0072 | 0.018 | 0.036 | 0.072 | 0.18 | 0.36 | 0.036
C12_10Cl | 0.002 | 0.00014 | 0.00035 | 0.0007 | 0.0014 | 0.0035 | 0.007 | 0.014 | 0.035 | 0.07 | 0.007
C13_04Cl | 0.005 | 0.00028 | 0.0007 | 0.0014 | 0.0028 | 0.007 | 0.014 | 0.028 | 0.07 | 0.14 | 0.014
C13_05Cl | 0.002 | 0.00078 | 0.00195 | 0.0039 | 0.0078 | 0.0195 | 0.039 | 0.078 | 0.195 | 0.39 | 0.039
C13_06Cl | 0.002 | 0.00094 | 0.00235 | 0.0047 | 0.0094 | 0.0235 | 0.047 | 0.094 | 0.235 | 0.47 | 0.047
C13_07Cl | 0.002 | 0.0008 | 0.002 | 0.004 | 0.008 | 0.02 | 0.04 | 0.08 | 0.2 | 0.4 | 0.04
C13_08Cl | 0.002 | 0.00104 | 0.0026 | 0.0052 | 0.0104 | 0.026 | 0.052 | 0.104 | 0.26 | 0.52 | 0.052
C13_09Cl | 0.002 | 0.00082 | 0.00205 | 0.0041 | 0.0082 | 0.0205 | 0.041 | 0.082 | 0.205 | 0.41 | 0.041
C13_10Cl | 0.002 | 0.0003 | 0.00075 | 0.0015 | 0.003 | 0.0075 | 0.015 | 0.03 | 0.075 | 0.15 | 0.015
D18-HBCD |  | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15
|  | 0.02 | 0.05 | 0.1 | 0.2 | 0.5 | 1 | 2 | 5 | 10 | 0.997
Responce factor | [10 empty] | Average | Coefficient variation
C10_04Cl | [10 empty] | #DIV/0! | #DIV/0!
C10_05Cl | [10 empty] | #DIV/0! | #DIV/0!
C10_06Cl | [10 empty] | #DIV/0! | #DIV/0!
C10_07Cl | [10 empty] | #DIV/0! | #DIV/0!
C10_08Cl | [10 empty] | #DIV/0! | #DIV/0!
C11_04Cl | [10 empty] | #DIV/0! | #DIV/0!
C11_05Cl | [10 empty] | #DIV/0! | #DIV/0!
C11_06Cl | [10 empty] | #DIV/0! | #DIV/0!
C11_07Cl | [10 empty] | #DIV/0! | #DIV/0!
C11_08Cl | [10 empty] | #DIV/0! | #DIV/0!
C11_09Cl | [10 empty] | #DIV/0! | #DIV/0!
C12_04Cl | [10 empty] | #DIV/0! | #DIV/0!
C12_05Cl | [10 empty] | #DIV/0! | #DIV/0!
C12_06Cl | [10 empty] | #DIV/0! | #DIV/0!
C12_07Cl | [10 empty] | #DIV/0! | #DIV/0!
C12_08Cl | [10 empty] | #DIV/0! | #DIV/0!
C12_09Cl | [10 empty] | #DIV/0! | #DIV/0!
C12_10Cl | [10 empty] | #DIV/0! | #DIV/0!
C13_04Cl | [10 empty] | #DIV/0! | #DIV/0!
C13_05Cl | [10 empty] | #DIV/0! | #DIV/0!
C13_06Cl | [10 empty] | #DIV/0! | #DIV/0!
C13_07Cl | [10 empty] | #DIV/0! | #DIV/0!
C13_08Cl | [10 empty] | #DIV/0! | #DIV/0!
C13_09Cl | [10 empty] | #DIV/0! | #DIV/0!
C13_10Cl | [10 empty] | #DIV/0! | #DIV/0!
|  | C10 | C11 | C12 | C13
| Cl4 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| Cl5 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| Cl6 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| Cl7 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| Cl8 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| Cl9 |  | #DIV/0! | #DIV/0! | #DIV/0!
| Cl10 |  |  | #DIV/0! | #DIV/0!
%Cl
C10_Cl4 | 50.6663 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.025333 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl5 | 56.3235 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.011265 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl6 | 60.8763 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.012175 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl7 | 64.6194 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.003231 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C10_Cl8 | 67.9052 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0013581 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl4 | 48.2512 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0096502 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl5 | 53.9167 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0026958 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl6 | 58.519 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0011704 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl7 | 62.3318 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0012466 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl8 | 65.7014 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.001314 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C11_Cl9 | 68.4183 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0013684 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl4 | 46.0558 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0046056 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl5 | 51.7071 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0025854 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl6 | 56.3375 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0011267 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl7 | 60.2006 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.001204 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl8 | 63.6362 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0012727 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl9 | 66.4199 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0013284 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C12_Cl10 | 68.8356 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0013767 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl4 | 44.0515 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0022026 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl5 | 49.6716 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.00099343 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl6 | 54.3128 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0010863 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl7 | 58.2104 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0011642 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl8 | 61.6969 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0012339 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl9 | 64.5349 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0012907 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C13_Cl10 | 67.0077 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0013402 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## MCCP cal_1
MCCP
[13 empty] | 220915_Blank_1 | 220915_1 | 220915_2 | 220915_3 | 220915_4 | 220915_5 | 220915_6 | 220915_7 | 220915_8 | 220915_9
Sample ID |  | STD1 | STD2 | STD3 | STD4 | STD5 | STD6 | STD7 | STD8 | STD9 |  |  | Sample01 | Sample02 | Sample03 | Sample04 | Sample05 | Sample06 | Sample07 | Sample08 | Sample09 | Sample10 | STD5
Sample name | IDL | MCCP-MIX-CS1 | MCCP-MIX-CS2 | MCCP-MIX-CS3 | MCCP-MIX-CS4 | MCCP-MIX-CS5 | MCCP-MIX-CS6 | MCCP-MIX-CS7 | MCCP-MIX-CS8 | MCCP-MIX-CS9 |  | MDL | 220915_9 STD5: MCCP-MIX-CS5
Sample | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 1
Extract-1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 1
Extract-2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 1
LC vial | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 1 | bar: Sample01 to 220915_9 STD5
D18-HBCD (μg/mL) | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 |  | 0.15 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.6 | 0.3 | 0.3 | 0.3 | 0.15
Unit | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL |  | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/mL
C14_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl4-10 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl4-10 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl4-10 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl4-10 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
MCCP |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
%Cl |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  | - | - | - | - | - | - | - | - | - | - | 0
Concentrations of target compounds in sample
Unit | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL |  | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/mL
C14_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Concentrations of target compounds in final solvent
Unit | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL |  | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL
C14_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Measurement result
C14_Cl4 | #DIV/0!
C14_Cl5 | #DIV/0!
C14_Cl6 | #DIV/0!
C14_Cl7 | #DIV/0!
C14_Cl8 | #DIV/0!
C14_Cl9 | #DIV/0!
C14_Cl10 | #DIV/0!
C15_Cl4 | #DIV/0!
C15_Cl5 | #DIV/0!
C15_Cl6 | #DIV/0!
C15_Cl7 | #DIV/0!
C15_Cl8 | #DIV/0!
C15_Cl9 | #DIV/0!
C15_Cl10 | #DIV/0!
C16_Cl4 | #DIV/0!
C16_Cl5 | #DIV/0!
C16_Cl6 | #DIV/0!
C16_Cl7 | #DIV/0!
C16_Cl8 | #DIV/0!
C16_Cl9 | #DIV/0!
C16_Cl10 | #DIV/0!
C17_Cl4 | #DIV/0!
C17_Cl5 | #DIV/0!
C17_Cl6 | #DIV/0!
C17_Cl7 | #DIV/0!
C17_Cl8 | #DIV/0!
C17_Cl9 | #DIV/0!
C17_Cl10 | #DIV/0!
D18-HBCD
Calibration solution concentration
Unit | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | Ratio
C14_Cl4 | 0.002 | 0.00198 | 0.00495 | 0.0099 | 0.0198 | 0.0495 | 0.099 | 0.198 | 0.495 | 0.99 | 0.099
C14_Cl5 | 0.002 | 0.0032 | 0.008 | 0.016 | 0.032 | 0.08 | 0.16 | 0.32 | 0.8 | 1.6 | 0.16
C14_Cl6 | 0.002 | 0.0028 | 0.007 | 0.014 | 0.028 | 0.07 | 0.14 | 0.28 | 0.7 | 1.4 | 0.14
C14_Cl7 | 0.002 | 0.0026 | 0.0065 | 0.013 | 0.026 | 0.065 | 0.13 | 0.26 | 0.65 | 1.3 | 0.13
C14_Cl8 | 0.002 | 0.0022 | 0.0055 | 0.011 | 0.022 | 0.055 | 0.11 | 0.22 | 0.55 | 1.1 | 0.11
C14_Cl9 | 0.002 | 0.0008 | 0.002 | 0.004 | 0.008 | 0.02 | 0.04 | 0.08 | 0.2 | 0.4 | 0.04
C14_Cl10 | 0.002 | 0.000138 | 0.000345 | 0.00069 | 0.00138 | 0.00345 | 0.0069 | 0.0138 | 0.0345 | 0.069 | 0.0069
C15_Cl4 | 0.002 | 0.0005 | 0.00125 | 0.0025 | 0.005 | 0.0125 | 0.025 | 0.05 | 0.125 | 0.25 | 0.025
C15_Cl5 | 0.002 | 0.00102 | 0.00255 | 0.0051 | 0.0102 | 0.0255 | 0.051 | 0.102 | 0.255 | 0.51 | 0.051
C15_Cl6 | 0.002 | 0.00094 | 0.00235 | 0.0047 | 0.0094 | 0.0235 | 0.047 | 0.094 | 0.235 | 0.47 | 0.047
C15_Cl7 | 0.002 | 0.00096 | 0.0024 | 0.0048 | 0.0096 | 0.024 | 0.048 | 0.096 | 0.24 | 0.48 | 0.048
C15_Cl8 | 0.002 | 0.00096 | 0.0024 | 0.0048 | 0.0096 | 0.024 | 0.048 | 0.096 | 0.24 | 0.48 | 0.048
C15_Cl9 | 0.002 | 0.00052 | 0.0013 | 0.0026 | 0.0052 | 0.013 | 0.026 | 0.052 | 0.13 | 0.26 | 0.026
C15_Cl10 | 0.002 | 0.000124 | 0.00031 | 0.00062 | 0.00124 | 0.0031 | 0.0062 | 0.0124 | 0.031 | 0.062 | 0.0062
C16_Cl4 | 0.002 | 5.8e-05 | 0.000145 | 0.00029 | 0.00058 | 0.00145 | 0.0029 | 0.0058 | 0.0145 | 0.029 | 0.0029
C16_Cl5 | 0.002 | 0.00014 | 0.00035 | 0.0007 | 0.0014 | 0.0035 | 0.007 | 0.014 | 0.035 | 0.07 | 0.007
C16_Cl6 | 0.002 | 0.000154 | 0.000385 | 0.00077 | 0.00154 | 0.00385 | 0.0077 | 0.0154 | 0.0385 | 0.077 | 0.0077
C16_Cl7 | 0.002 | 0.000194 | 0.000485 | 0.00097 | 0.00194 | 0.00485 | 0.0097 | 0.0194 | 0.0485 | 0.097 | 0.0097
C16_Cl8 | 0.002 | 0.00026 | 0.00065 | 0.0013 | 0.0026 | 0.0065 | 0.013 | 0.026 | 0.065 | 0.13 | 0.013
C16_Cl9 | 0.002 | 0.000194 | 0.000485 | 0.00097 | 0.00194 | 0.00485 | 0.0097 | 0.0194 | 0.0485 | 0.097 | 0.0097
C16_Cl10 | 0.002 | 6.2e-05 | 0.000155 | 0.00031 | 0.00062 | 0.00155 | 0.0031 | 0.0062 | 0.0155 | 0.031 | 0.0031
C17_Cl4 | 0.002 | 6.2e-06 | 1.55e-05 | 3.1e-05 | 6.2e-05 | 0.000155 | 0.00031 | 0.00062 | 0.00155 | 0.0031 | 0.00031
C17_Cl5 | 0.002 | 1.96e-05 | 4.9e-05 | 9.8e-05 | 0.000196 | 0.00049 | 0.00098 | 0.00196 | 0.0049 | 0.0098 | 0.00098
C17_Cl6 | 0.002 | 1.96e-05 | 4.9e-05 | 9.8e-05 | 0.000196 | 0.00049 | 0.00098 | 0.00196 | 0.0049 | 0.0098 | 0.00098
C17_Cl7 | 0.002 | 4.2e-05 | 0.000105 | 0.00021 | 0.00042 | 0.00105 | 0.0021 | 0.0042 | 0.0105 | 0.021 | 0.0021
C17_Cl8 | 0.002 | 7.8e-05 | 0.000195 | 0.00039 | 0.00078 | 0.00195 | 0.0039 | 0.0078 | 0.0195 | 0.039 | 0.0039
C17_Cl9 | 0.002 | 7.2e-05 | 0.00018 | 0.00036 | 0.00072 | 0.0018 | 0.0036 | 0.0072 | 0.018 | 0.036 | 0.0036
C17_Cl10 | 0.002 | 2.6e-05 | 6.5e-05 | 0.00013 | 0.00026 | 0.00065 | 0.0013 | 0.0026 | 0.0065 | 0.013 | 0.0013
D18-HBCD |  | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15
|  | 0.02 | 0.05 | 0.1 | 0.2 | 0.5 | 1 | 2 | 5 | 10 | 1.0034
Responce factor | [10 empty] | Average | Coefficient variation
C14_Cl4 | [10 empty] | #DIV/0! | #DIV/0!
C14_Cl5 | [10 empty] | #DIV/0! | #DIV/0!
C14_Cl6 | [10 empty] | #DIV/0! | #DIV/0!
C14_Cl7 | [10 empty] | #DIV/0! | #DIV/0!
C14_Cl8 | [10 empty] | #DIV/0! | #DIV/0!
C14_Cl9 | [10 empty] | #DIV/0! | #DIV/0!
C14_Cl10 | [10 empty] | #DIV/0! | #DIV/0!
C15_Cl4 | [10 empty] | #DIV/0! | #DIV/0!
C15_Cl5 | [10 empty] | #DIV/0! | #DIV/0!
C15_Cl6 | [10 empty] | #DIV/0! | #DIV/0!
C15_Cl7 | [10 empty] | #DIV/0! | #DIV/0!
C15_Cl8 | [10 empty] | #DIV/0! | #DIV/0!
C15_Cl9 | [10 empty] | #DIV/0! | #DIV/0!
C15_Cl10 | [10 empty] | #DIV/0! | #DIV/0!
C16_Cl4 | [10 empty] | #DIV/0! | #DIV/0!
C16_Cl5 | [10 empty] | #DIV/0! | #DIV/0!
C16_Cl6 | [10 empty] | #DIV/0! | #DIV/0!
C16_Cl7 | [10 empty] | #DIV/0! | #DIV/0!
C16_Cl8 | [10 empty] | #DIV/0! | #DIV/0!
C16_Cl9 | [10 empty] | #DIV/0! | #DIV/0!
C16_Cl10 | [10 empty] | #DIV/0! | #DIV/0!
C17_Cl4 | [10 empty] | #DIV/0! | #DIV/0!
C17_Cl5 | [10 empty] | #DIV/0! | #DIV/0!
C17_Cl6 | [10 empty] | #DIV/0! | #DIV/0!
C17_Cl7 | [10 empty] | #DIV/0! | #DIV/0!
C17_Cl8 | [10 empty] | #DIV/0! | #DIV/0!
C17_Cl9 | [10 empty] | #DIV/0! | #DIV/0!
C17_Cl10 | [10 empty] | #DIV/0! | #DIV/0!
|  | C14 | C15 | C16 | C17
| Cl4 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| Cl5 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| Cl6 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| Cl7 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| Cl8 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| Cl9 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
| Cl10 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
%Cl
C14_Cl4 | 41.869 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.00083738 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl5 | 47.5069 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.00095014 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl6 | 52.1922 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0010438 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl7 | 56.1476 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.001123 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl8 | 59.5312 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0011906 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl9 | 62.4588 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0012492 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C14_Cl10 | 65.0166 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0013003 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl4 | 40.1832 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.00080366 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl5 | 45.7641 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.00091528 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl6 | 50.4338 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0010087 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl7 | 54.3987 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.001088 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl8 | 57.8071 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0011561 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl9 | 60.7685 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0012154 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C15_Cl10 | 63.3654 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0012673 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl4 | 38.6279 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.00077256 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl5 | 44.1447 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.00088289 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl6 | 48.7901 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0009758 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl7 | 52.7555 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0010551 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl8 | 56.18 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0011236 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl9 | 59.1673 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0011833 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C16_Cl10 | 61.7959 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0012359 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl4 | 37.1885 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.00074377 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl5 | 42.6359 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.00085272 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl6 | 47.2501 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.000945 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl7 | 51.2087 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0010242 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl8 | 54.642 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.0010928 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl9 | 57.6483 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.001153 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C17_Cl10 | 60.3024 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.001206 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## LCCP calc_1
LCCP
[13 empty] | 220915_Blank_1 | 220915_1 | 220915_2 | 220915_3 | 220915_4 | 220915_5 | 220915_6 | 220915_7 | 220915_8 | 220915_9
Sample ID |  | STD1 | STD2 | STD3 | STD4 | STD5 | STD6 | STD7 | STD8 | STD9 |  |  | Sample01 | Sample02 | Sample03 | Sample04 | Sample05 | Sample06 | Sample07 | Sample08 | Sample09 | Sample10 | STD5
Sample name | IDL | LCCP-MIX-CS1 | LCCP-MIX-CS2 | LCCP-MIX-CS3 | LCCP-MIX-CS4 | LCCP-MIX-CS5 | LCCP-MIX-CS6 | LCCP-MIX-CS7 | LCCP-MIX-CS8 | LCCP-MIX-CS9 |  | MDL | 220915_9 STD5: LCCP-MIX-CS5
Sample | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 1
Extract-1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 1
Extract-2 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 1
LC vial | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | 1 | bar: Sample01 to 220915_9 STD5
D18-HBCD (μg/mL) | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 |  | 0.15 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.15
Unit | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL |  | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/mL
C18_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl4-10 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl4-10 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl4-10 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
LCCP |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
%Cl |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  | - | - | - | - | - | - | - | - | - | - | 0
Concentrations in samples
Unit | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL |  | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/g | μg/mL
C18_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Concentrations in LC vials
Unit | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL |  | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL
C18_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl4 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl5 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl6 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl7 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl8 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl9 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl10 | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0.002 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Abundances
ID
C18_Cl4 | #DIV/0!
C18_Cl5 | #DIV/0!
C18_Cl6 | #DIV/0!
C18_Cl7 | #DIV/0!
C18_Cl8 | #DIV/0!
C18_Cl9 | #DIV/0!
C18_Cl10 | #DIV/0!
C19_Cl4 | #DIV/0!
C19_Cl5 | #DIV/0!
C19_Cl6 | #DIV/0!
C19_Cl7 | #DIV/0!
C19_Cl8 | #DIV/0!
C19_Cl9 | #DIV/0!
C19_Cl10 | #DIV/0!
C20_Cl4 | #DIV/0!
C20_Cl5 | #DIV/0!
C20_Cl6 | #DIV/0!
C20_Cl7 | #DIV/0!
C20_Cl8 | #DIV/0!
C20_Cl9 | #DIV/0!
C20_Cl10 | #DIV/0!
D18_HBCD
Concentrations in calibration solutions
Unit | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | μg/mL | Ratio
C18_Cl4 | 0.002 | 0.0012585 | 0.0031462 | 0.0062925 | 0.012585 | 0.031462 | 0.062925 | 0.12585 | 0.31462 | 0.62925 | 0.062925 | 0.12296 | Sample01: 0.0028875; 220915_1 Sample02: 0.062925
C18_Cl5 | 0.002 | 0.0021407 | 0.0053518 | 0.010704 | 0.021407 | 0.053518 | 0.10704 | 0.21407 | 0.53518 | 1.0704 | 0.10704 | 0.19149 | Sample01: 0.022584; 220915_1 Sample02: 0.10704
C18_Cl6 | 0.002 | 0.0024074 | 0.0060185 | 0.012037 | 0.024074 | 0.060185 | 0.12037 | 0.24074 | 0.60185 | 1.2037 | 0.12037 | 0.15022 | Sample01: 0.090519; 220915_1 Sample02: 0.12037
C18_Cl7 | 0.002 | 0.0021247 | 0.0053118 | 0.010624 | 0.021247 | 0.053118 | 0.10624 | 0.21247 | 0.53118 | 1.0624 | 0.10624 | 0.056821 | Sample01: 0.15565; 220915_1 Sample02: 0.10624
C18_Cl8 | 0.002 | 0.0014025 | 0.0035064 | 0.0070127 | 0.014025 | 0.035064 | 0.070127 | 0.14025 | 0.35064 | 0.70127 | 0.070127 | 0.011362 | Sample01: 0.12889; 220915_1 Sample02: 0.070127
C18_Cl9 | 0.002 | 0.00054034 | 0.0013508 | 0.0027017 | 0.0054034 | 0.013508 | 0.027017 | 0.054034 | 0.13508 | 0.27017 | 0.027017 | 0.0017242 | Sample01: 0.05231; 220915_1 Sample02: 0.027017
C18_Cl10 | 0.002 | 0.0001076 | 0.000269 | 0.00053799 | 0.001076 | 0.00269 | 0.0053799 | 0.01076 | 0.0269 | 0.053799 | 0.0053799 | 0.00021287 | Sample01: 0.010547; 220915_1 Sample02: 0.0053799
C19_Cl4 | 0.002 | 0.00033238 | 0.00083095 | 0.0016619 | 0.0033238 | 0.0083095 | 0.016619 | 0.033238 | 0.083095 | 0.16619 | 0.016619 | 0.032702 | Sample01: 0.00053611; 220915_1 Sample02: 0.016619
C19_Cl5 | 0.002 | 0.00063092 | 0.0015773 | 0.0031546 | 0.0063092 | 0.015773 | 0.031546 | 0.063092 | 0.15773 | 0.31546 | 0.031546 | 0.058064 | Sample01: 0.0050278; 220915_1 Sample02: 0.031546
C19_Cl6 | 0.002 | 0.00076842 | 0.0019211 | 0.0038421 | 0.0076842 | 0.019211 | 0.038421 | 0.076842 | 0.19211 | 0.38421 | 0.038421 | 0.053543 | Sample01: 0.0233; 220915_1 Sample02: 0.038421
C19_Cl7 | 0.002 | 0.00072645 | 0.0018161 | 0.0036322 | 0.0072645 | 0.018161 | 0.036322 | 0.072645 | 0.18161 | 0.36322 | 0.036322 | 0.022673 | Sample01: 0.049971; 220915_1 Sample02: 0.036322
C19_Cl8 | 0.002 | 0.00055677 | 0.0013919 | 0.0027839 | 0.0055677 | 0.013919 | 0.027839 | 0.055677 | 0.13919 | 0.27839 | 0.027839 | 0.0050994 | Sample01: 0.050578; 220915_1 Sample02: 0.027839
C19_Cl9 | 0.002 | 0.00024904 | 0.00062259 | 0.0012452 | 0.0024904 | 0.0062259 | 0.012452 | 0.024904 | 0.062259 | 0.12452 | 0.012452 | 0.00080203 | Sample01: 0.024101; 220915_1 Sample02: 0.012452
C19_Cl10 | 0.002 | 5.9627e-05 | 0.00014907 | 0.00029814 | 0.00059627 | 0.0014907 | 0.0029814 | 0.0059627 | 0.014907 | 0.029814 | 0.0029814 | 9.4457e-05 | Sample01: 0.0058683; 220915_1 Sample02: 0.0029814
C20_Cl4 | 0.002 | 6.5803e-05 | 0.00016451 | 0.00032901 | 0.00065803 | 0.0016451 | 0.0032901 | 0.0065803 | 0.016451 | 0.032901 | 0.0032901 | 0.0065803 | Sample01: 0; 220915_1 Sample02: 0.0032901
C20_Cl5 | 0.002 | 0.00014076 | 0.0003519 | 0.0007038 | 0.0014076 | 0.003519 | 0.007038 | 0.014076 | 0.03519 | 0.07038 | 0.007038 | 0.013553 | Sample01: 0.00052352; 220915_1 Sample02: 0.007038
C20_Cl6 | 0.002 | 0.00018636 | 0.0004659 | 0.00093179 | 0.0018636 | 0.004659 | 0.0093179 | 0.018636 | 0.04659 | 0.093179 | 0.0093179 | 0.014635 | Sample01: 0.004001; 220915_1 Sample02: 0.0093179
C20_Cl7 | 0.002 | 0.00017239 | 0.00043098 | 0.00086196 | 0.0017239 | 0.0043098 | 0.0086196 | 0.017239 | 0.043098 | 0.086196 | 0.0086196 | 0.0069921 | Sample01: 0.010247; 220915_1 Sample02: 0.0086196
C20_Cl8 | 0.002 | 0.0001373 | 0.00034325 | 0.00068651 | 0.001373 | 0.0034325 | 0.0068651 | 0.01373 | 0.034325 | 0.068651 | 0.0068651 | 0.0018376 | Sample01: 0.011892; 220915_1 Sample02: 0.0068651
C20_Cl9 | 0.002 | 6.8635e-05 | 0.00017159 | 0.00034318 | 0.00068635 | 0.0017159 | 0.0034318 | 0.0068635 | 0.017159 | 0.034318 | 0.0034318 | 0.00039648 | Sample01: 0.0064671; 220915_1 Sample02: 0.0034318
C20_Cl10 | 0.002 | 1.7677e-05 | 4.4194e-05 | 8.8387e-05 | 0.00017677 | 0.00044194 | 0.00088387 | 0.0017677 | 0.0044194 | 0.0088387 | 0.00088387 | 0 | Sample01: 0.0017677; 220915_1 Sample02: 0.00088387
D18-HBCD |  | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15
|  | 0.02 | 0.05 | 0.1 | 0.2 | 0.5 | 1 | 2 | 5 | 10 | 0.70472 |  | 220915_1 Sample02: 0.70472
Responce factors | [10 empty] | Average | Coefficient variation
C18_Cl4 | [10 empty] | #DIV/0! | #DIV/0!
C18_Cl5 | [10 empty] | #DIV/0! | #DIV/0!
C18_Cl6 | [10 empty] | #DIV/0! | #DIV/0!
C18_Cl7 | [10 empty] | #DIV/0! | #DIV/0!
C18_Cl8 | [10 empty] | #DIV/0! | #DIV/0!
C18_Cl9 | [10 empty] | #DIV/0! | #DIV/0!
C18_Cl10 | [10 empty] | #DIV/0! | #DIV/0!
C19_Cl4 | [10 empty] | #DIV/0! | #DIV/0!
C19_Cl5 | [10 empty] | #DIV/0! | #DIV/0!
C19_Cl6 | [10 empty] | #DIV/0! | #DIV/0!
C19_Cl7 | [10 empty] | #DIV/0! | #DIV/0!
C19_Cl8 | [10 empty] | #DIV/0! | #DIV/0!
C19_Cl9 | [10 empty] | #DIV/0! | #DIV/0!
C19_Cl10 | [10 empty] | #DIV/0! | #DIV/0!
C20_Cl4 | [10 empty] | #DIV/0! | #DIV/0!
C20_Cl5 | [10 empty] | #DIV/0! | #DIV/0!
C20_Cl6 | [10 empty] | #DIV/0! | #DIV/0!
C20_Cl7 | [10 empty] | #DIV/0! | #DIV/0!
C20_Cl8 | [10 empty] | #DIV/0! | #DIV/0!
C20_Cl9 | [10 empty] | #DIV/0! | #DIV/0!
C20_Cl10 | [10 empty] | #DIV/0! | #DIV/0!
|  | C18 | C19 | C20
| Cl4 | #DIV/0! | #DIV/0! | #DIV/0!
| Cl5 | #DIV/0! | #DIV/0! | #DIV/0!
| Cl6 | #DIV/0! | #DIV/0! | #DIV/0!
| Cl7 | #DIV/0! | #DIV/0! | #DIV/0!
| Cl8 | #DIV/0! | #DIV/0! | #DIV/0!
| Cl9 | #DIV/0! | #DIV/0! | #DIV/0!
| Cl10 | #DIV/0! | #DIV/0! | #DIV/0!
%Cl
C18_Cl4 | 41.869 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.00083738 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl5 | 47.5069 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.00095014 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl6 | 52.1922 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.0010438 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl7 | 56.1476 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.001123 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl8 | 59.5312 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.0011906 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl9 | 62.4588 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.0012492 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C18_Cl10 | 65.0166 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.0013003 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl4 | 40.1832 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.00080366 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl5 | 45.7641 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.00091528 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl6 | 50.4338 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.0010087 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl7 | 54.3987 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.001088 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl8 | 57.8071 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.0011561 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl9 | 60.7685 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.0012154 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C19_Cl10 | 63.3654 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.0012673 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl4 | 38.6279 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.00077256 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl5 | 44.1447 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.00088289 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl6 | 48.7901 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.0009758 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl7 | 52.7555 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.0010551 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl8 | 56.18 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.0011236 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl9 | 59.1673 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.0011833 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
C20_Cl10 | 61.7959 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  |  |  | 0.0012359 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## SCCP method
SCCP測定条件
UPLC system | Waters ACQUITY UPLC H-Class
Column | Agilent ZORBAX SB-CN 2.1x100 mm, 1.8 μm
Column Temp. | 40 °C
Mobile phase A | 5 mM CH3COONH4 water
Mobile phase B | 5 mM CH3COONH4 MeOH
Injection volume | 5 μL
Gradient
| Time (min) | %B | Flow rate (mL/min)
| 0 | 60 | 0.4
| 5 | 73 | 0.4
| 20 | 99 | 0.4
| 22 | 99 | 0.4
| 22.1 | 60 | 0.4
| 25 | 60 | 0.4
Tandem MS system | Waters Xevo TQ-S micro
Ionization | ESI negative
Gas Temp | 110 °C
Drying Gas | 10 L/min
Nebulizer | 50 psig
Sheath Gas Temp | 110 °C
Sheath Gas Flow | 6 L/min
Vcap | 3000 V
Nozzle Voltage
Fragmentor | 100 V
Skimmer | 65 V
OCT 1RF Vpp | 750 V
| MRM1 |  |  |  | MRM2
| Precursor ion (m/z) | Cone voltage (V) | Collision energy (eV) | Product ion (m/z) | Precursor ion (m/z) | Cone voltage (V) | Collision energy (eV) | Product ion (m/z)
C10Cl4 | 339.0272 | 20 | 10 | 59.0133 | 337.0301 | 20 | 10 | 59.0133 | m/z 339>59
C10Cl5 | 372.9882 | 20 | 20 | 59.0133 | 370.9911 | 20 | 20 | 59.0133 | m/z 373>59
C10Cl6 | 406.9492 | 20 | 5 | 346.9275 | 404.9522 | 20 | 5 | 344.9305 | m/z 407>347
C10Cl7 | 440.9102 | 20 | 5 | 380.8886 | 438.9132 | 20 | 5 | 378.8915 | m/z 441>381
C10Cl8 | 476.8683 | 20 | 5 | 416.8466 | 474.8713 | 20 | 5 | 414.8496 | m/z 477>417
C11Cl4 | 353.0428 | 20 | 10 | 59.0133 | 351.0458 | 20 | 10 | 59.0133 | m/z 353>59
C11Cl5 | 387.0038 | 20 | 25 | 59.0133 | 385.0068 | 20 | 25 | 59.0133 | m/z 387>59
C11Cl6 | 420.9649 | 20 | 5 | 360.9432 | 418.9678 | 20 | 5 | 358.9461 | m/z 421>361
C11Cl7 | 454.9259 | 20 | 5 | 394.9042 | 452.9288 | 20 | 5 | 392.9072 | m/z 455>395
C11Cl8 | 490.884 | 20 | 5 | 430.8623 | 488.8869 | 20 | 5 | 428.8652 | m/z 491>431
C11Cl9 | 524.845 | 20 | 5 | 464.8233 | 522.848 | 20 | 5 | 462.8263 | m/z 525>465
C12Cl4 | 367.0585 | 20 | 10 | 59.0133 | 365.0614 | 20 | 10 | 59.0133 | m/z 367>59
C12Cl5 | 401.0195 | 20 | 25 | 59.0133 | 399.0224 | 20 | 25 | 59.0133 | m/z 401>59
C12Cl6 | 434.9805 | 20 | 5 | 374.9588 | 432.9835 | 20 | 5 | 372.9618 | m/z 435>375
C12Cl7 | 468.9415 | 20 | 5 | 408.9199 | 466.9445 | 20 | 5 | 406.9228 | m/z 469>409
C12Cl8 | 504.8996 | 20 | 5 | 444.8779 | 502.9026 | 20 | 5 | 442.8809 | m/z 505>445
C12Cl9 | 538.8607 | 20 | 5 | 478.839 | 536.8636 | 20 | 5 | 476.8419 | m/z 539>479
C12Cl10 | 572.8217 | 20 | 5 | 512.8 | 570.8246 | 20 | 5 | 510.803 | m/z 573>513
C13Cl4 | 381.0741 | 20 | 15 | 59.0133 | 379.0771 | 20 | 15 | 59.0133 | m/z 381>59
C13Cl5 | 415.0351 | 20 | 25 | 59.0133 | 413.0381 | 20 | 25 | 59.0133 | m/z 415>59
C13Cl6 | 448.9962 | 20 | 5 | 388.9745 | 446.9991 | 20 | 5 | 386.9774 | m/z 449>389
C13Cl7 | 482.9572 | 20 | 5 | 422.9355 | 480.9601 | 20 | 5 | 420.9385 | m/z 483>423
C13Cl8 | 518.9153 | 20 | 5 | 458.8936 | 516.9182 | 20 | 5 | 456.8965 | m/z 519>459
C13Cl9 | 552.8763 | 20 | 5 | 492.8546 | 550.8793 | 20 | 5 | 490.8576 | m/z 553>493
C13Cl10 | 586.8373 | 20 | 5 | 526.8157 | 584.8403 | 20 | 5 | 524.8186 | m/z 587>527

## SCCP chromatogram
SCCPクロマトグラム

## MCCP method
MCCP測定条件
UPLC system | Waters ACQUITY UPLC H-Class
Column | Agilent ZORBAX SB-CN 2.1x100 mm, 1.8 μm
Column Temp. | 40 °C
Mobile phase A | 5 mM CH3COONH4 water
Mobile phase B | 5 mM CH3COONH4 MeOH
Injection volume | 5 μL
Gradient
| Time (min) | %B | Flow rate (mL/min)
| 0 | 60 | 0.4
| 5 | 73 | 0.4
| 20 | 99 | 0.4
| 22 | 99 | 0.4
| 22.1 | 60 | 0.4
| 25 | 60 | 0.4
Tandem MS system | Waters Xevo TQ-S micro
Ionization | ESI negative
Capillary voltage | 0.75 kV
Source temperature | 110 °C
Desolvation temperature | 500 °C
Desolvation gas flow | 1000 L/hr
Cone gas flow | 20 L/hr
Cone voltage | 20 V
| MRM1 |  |  |  | MRM2
| Precursor ion (m/z) | Cone voltage (V) | Collision energy (eV) | Product ion (m/z) | Precursor ion (m/z) | Cone voltage (V) | Collision energy (eV) | Product ion (m/z)
C14Cl4 | 395.0892 | 20 | 10 | 59.0133 | 393.0922 | 20 | 10 | 59.0133 | m/z 395>59
C14Cl5 | 429.0502 | 20 | 15 | 59.0133 | 427.0532 | 20 | 15 | 59.0133 | m/z 429>59
C14Cl6 | 463.0113 | 20 | 20 | 59.0133 | 461.0142 | 20 | 20 | 59.0133 | m/z 463>59
C14Cl7 | 496.9723 | 20 | 5 | 436.9512 | 494.9753 | 20 | 5 | 434.9541 | m/z 497>437
C14Cl8 | 532.9304 | 20 | 5 | 472.9092 | 530.9333 | 20 | 5 | 470.9122 | m/z 533>473
C14Cl9 | 566.8914 | 20 | 5 | 506.8703 | 564.8944 | 20 | 5 | 504.8732 | m/z 567>507
C14Cl10 | 600.8524 | 20 | 5 | 540.8313 | 598.8554 | 20 | 5 | 538.8343 | m/z 601>541
C15Cl4 | 409.1049 | 20 | 10 | 59.0133 | 407.1078 | 20 | 10 | 59.0133 | m/z 409>59
C15Cl5 | 443.0659 | 20 | 20 | 59.0133 | 441.0688 | 20 | 20 | 59.0133 | m/z 443>59
C15Cl6 | 477.0269 | 20 | 20 | 59.0133 | 475.0299 | 20 | 20 | 59.0133 | m/z 477>59
C15Cl7 | 510.988 | 20 | 5 | 450.9668 | 508.9909 | 20 | 5 | 448.9698 | m/z 511>451
C15Cl8 | 546.946 | 20 | 10 | 486.9249 | 544.949 | 20 | 10 | 484.9278 | m/z 547>487
C15Cl9 | 580.9071 | 20 | 10 | 520.8859 | 578.91 | 20 | 10 | 518.8889 | m/z 581>521
C15Cl10 | 614.8681 | 20 | 10 | 554.847 | 612.871 | 20 | 10 | 552.8499 | m/z 615>555
C16Cl4 | 423.1205 | 20 | 10 | 59.0133 | 421.1235 | 20 | 10 | 59.0133 | m/z 423>59
C16Cl5 | 457.0815 | 20 | 20 | 59.0133 | 455.0845 | 20 | 20 | 59.0133 | m/z 457>59
C16Cl6 | 491.0426 | 20 | 20 | 59.0133 | 489.0455 | 20 | 20 | 59.0133 | m/z 491>59
C16Cl7 | 525.0036 | 20 | 10 | 464.9825 | 523.0066 | 20 | 10 | 462.9854 | m/z 525>465
C16Cl8 | 560.9617 | 20 | 10 | 500.9405 | 558.9646 | 20 | 10 | 498.9435 | m/z 561>501
C16Cl9 | 594.9227 | 20 | 10 | 534.9016 | 592.9257 | 20 | 10 | 532.9045 | m/z 595>535
C16Cl10 | 628.8837 | 20 | 10 | 568.8626 | 626.8867 | 20 | 10 | 566.8656 | m/z 629>569
C17Cl4 | 437.1362 | 20 | 10 | 59.0133 | 435.1391 | 20 | 10 | 59.0133 | m/z 437>59
C17Cl5 | 471.0972 | 20 | 20 | 59.0133 | 469.1001 | 20 | 20 | 59.0133 | m/z 471>59
C17Cl6 | 505.0582 | 20 | 20 | 59.0133 | 503.0612 | 20 | 20 | 59.0133 | m/z 505>59
C17Cl7 | 539.0193 | 20 | 10 | 478.9981 | 537.0222 | 20 | 10 | 477.0011 | m/z 539>479
C17Cl8 | 574.9773 | 20 | 10 | 514.9562 | 572.9803 | 20 | 10 | 512.9591 | m/z 575>515
C17Cl9 | 608.9384 | 20 | 10 | 548.9172 | 606.9413 | 20 | 10 | 546.9202 | m/z 609>549
C17Cl10 | 642.8994 | 20 | 10 | 582.8783 | 640.9023 | 20 | 10 | 580.8812 | m/z 643>583

## MCCP chromatogram
MCCPクロマトグラム

## LCCP method
LCCP測定条件
UPLC system | Waters ACQUITY UPLC H-Class
Column | Agilent ZORBAX SB-CN 2.1x100 mm, 1.8 μm
Column Temp. | 40 °C
Mobile phase A | 5 mM CH3COONH4 water
Mobile phase B | 5 mM CH3COONH4 MeOH
Injection volume | 5 μL
Gradient
| Time (min) | %B | Flow rate (mL/min)
| 0 | 60 | 0.4
| 5 | 73 | 0.4
| 20 | 99 | 0.4
| 22 | 99 | 0.4
| 22.1 | 60 | 0.4
| 25 | 60 | 0.4
Tandem MS system | Waters Xevo TQ-S micro
Ionization | ESI negative
Capillary voltage | 0.75 kV
Source temperature | 110 °C
Desolvation temperature | 500 °C
Desolvation gas flow | 1000 L/hr
Cone gas flow | 20 L/hr
Cone voltage | 20 V
| MRM1 |  |  |  | MRM2
| Precursor ion (m/z) | Cone voltage (V) | Collision energy (eV) | Product ion (m/z) | Precursor ion (m/z) | Cone voltage (V) | Collision energy (eV) | Product ion (m/z)
C18Cl4 | 451.1518 | 20 | 20 | 59.0133 | 449.1548 | 20 | 20 | 59.0133
C18Cl5 | 485.1128 | 20 | 20 | 59.0133 | 483.1158 | 20 | 20 | 59.0133
C18Cl6 | 519.0739 | 20 | 20 | 59.0133 | 517.0768 | 20 | 20 | 59.0133
C18Cl7 | 553.0349 | 20 | 10 | 493.0138 | 551.0379 | 20 | 10 | 491.0167
C18Cl8 | 588.993 | 20 | 10 | 528.9718 | 586.9959 | 20 | 10 | 526.9748
C18Cl9 | 622.954 | 20 | 10 | 562.9329 | 620.957 | 20 | 10 | 560.9358
C18Cl10 | 656.915 | 20 | 10 | 596.8939 | 654.918 | 20 | 10 | 594.8969
C19Cl4 | 465.1675 | 20 | 20 | 59.0133 | 463.1704 | 20 | 20 | 59.0133
C19Cl5 | 499.1285 | 20 | 20 | 59.0133 | 497.1314 | 20 | 20 | 59.0133
C19Cl6 | 533.0895 | 20 | 20 | 59.0133 | 531.0925 | 20 | 20 | 59.0133
C19Cl7 | 567.0506 | 20 | 10 | 507.0294 | 565.0535 | 20 | 10 | 505.0324
C19Cl8 | 603.0086 | 20 | 10 | 542.9875 | 601.0116 | 20 | 10 | 540.9904
C19Cl9 | 636.9697 | 20 | 10 | 576.9485 | 634.9726 | 20 | 10 | 574.9515
C19Cl10 | 670.9307 | 20 | 10 | 610.9096 | 668.9336 | 20 | 10 | 608.9125
C20Cl4 | 479.1831 | 20 | 20 | 59.0133 | 477.1861 | 20 | 20 | 59.0133
C20Cl5 | 513.1441 | 20 | 20 | 59.0133 | 511.1471 | 20 | 20 | 59.0133
C20Cl6 | 547.1052 | 20 | 20 | 59.0133 | 545.1081 | 20 | 20 | 59.0133
C20Cl7 | 581.0662 | 20 | 10 | 521.0451 | 579.0692 | 20 | 10 | 519.048
C20Cl8 | 617.0243 | 20 | 10 | 557.0031 | 615.0272 | 20 | 10 | 555.0061
C20Cl9 | 650.9853 | 20 | 10 | 590.9642 | 648.9883 | 20 | 10 | 588.9671
C20Cl10 | 684.9463 | 20 | 10 | 624.9252 | 682.9493 | 20 | 10 | 622.9282

## LCCP chromatogram
LCCPクロマトグラム

## TOF extract ions
TOF抽出イオン
|  | Molecular | Molecular | EIC m/z | EIC m/z | 1H | 2H | 12C | 13C | 16O | 35Cl | 37Cl | 79Br | 81Br | CH3COO | e | Time | SCCP-CS |  | MCCP-CS |  | LCCP-CS |  | Diable 700X
| Compound | formula | weight | [M-H]- | [M+CH3COO]- | 1.0078 | 2.0141 | 12 | 13.0034 | 15.9949 | 34.9689 | 36.9659 | 78.9183 | 80.9163 | 59.0133 | 0.00054858 | (min) | [M-H]- | [M+CH3COO]- | [M-H]- | [M+CH3COO]- | [M-H]- | [M+CH3COO]- | [M-H]- | [M+CH3COO]-
IS | HBCD-d18 | C12D18Br6 | 659.7577 | 657.7442 | 718.7716 |  | 18 | 12 |  |  |  |  | 3 | 3 |  |  | 12
vSCCP | C8Cl4 | C8H14Cl4 | 251.982 | 250.9747 | 310.9959 | 14 |  | 8 |  |  | 3 | 1
| C8Cl5 | C8H13Cl5 | 285.943 | 284.9358 | 344.9569 | 13 |  | 8 |  |  | 4 | 1
| C8Cl6 | C8H12Cl6 | 319.9041 | 318.8968 | 378.9179 | 12 |  | 8 |  |  | 5 | 1
| C8Cl7 | C8H11Cl7 | 353.8651 | 352.8578 | 412.8789 | 11 |  | 8 |  |  | 6 | 1
| C8Cl8 | C8H10Cl8 | 387.8261 | 386.8188 | 446.84 | 10 |  | 8 |  |  | 7 | 1
| C9Cl4 | C9H16Cl4 | 265.9977 | 264.9904 | 325.0115 | 16 |  | 9 |  |  | 3 | 1
| C9Cl5 | C9H15Cl5 | 299.9587 | 298.9514 | 358.9725 | 15 |  | 9 |  |  | 4 | 1
| C9Cl6 | C9H14Cl6 | 333.9197 | 332.9124 | 392.9336 | 14 |  | 9 |  |  | 5 | 1
| C9Cl7 | C9H13Cl7 | 367.8807 | 366.8735 | 426.8946 | 13 |  | 9 |  |  | 6 | 1
| C9Cl8 | C9H12Cl8 | 401.8418 | 400.8345 | 460.8556 | 12 |  | 9 |  |  | 7 | 1
| C9Cl9 | C9H11Cl9 | 435.8028 | 434.7955 | 494.8167 | 11 |  | 9 |  |  | 8 | 1
SCCP | C10Cl4 | C10H18Cl4 | 280.0133 | 279.006 | 339.0272 | 18 |  | 10 |  |  | 3 | 1 |  |  |  |  | 8.5
| C10Cl5 | C10H17Cl5 | 313.9743 | 312.9671 | 372.9882 | 17 |  | 10 |  |  | 4 | 1 |  |  |  |  | 9
| C10Cl6 | C10H16Cl6 | 347.9354 | 346.9281 | 406.9492 | 16 |  | 10 |  |  | 5 | 1 |  |  |  |  | 9.59
| C10Cl7 | C10H15Cl7 | 381.8964 | 380.8891 | 440.9102 | 15 |  | 10 |  |  | 6 | 1 |  |  |  |  | 10.2
| C10Cl8 | C10H14Cl8 | 415.8574 | 414.8501 | 474.8713 | 14 |  | 10 |  |  | 7 | 1 |  |  |  |  | 10.97
| C10Cl9 | C10H13Cl9 | 449.8184 | 448.8112 | 508.8323 | 13 |  | 10 |  |  | 8 | 1
| C10Cl10 | C10H12Cl10 | 483.7795 | 482.7722 | 542.7933 | 12 |  | 10 |  |  | 9 | 1
| C11Cl4 | C11H20Cl4 | 294.029 | 293.0217 | 353.0428 | 20 |  | 11 |  |  | 3 | 1 |  |  |  |  | 9.33
| C11Cl5 | C11H19Cl5 | 327.99 | 326.9827 | 387.0038 | 19 |  | 11 |  |  | 4 | 1 |  |  |  |  | 9.91
| C11Cl6 | C11H18Cl6 | 361.951 | 360.9437 | 420.9649 | 18 |  | 11 |  |  | 5 | 1 |  |  |  |  | 10.5
| C11Cl7 | C11H17Cl7 | 395.912 | 394.9048 | 454.9259 | 17 |  | 11 |  |  | 6 | 1 |  |  |  |  | 11.1
| C11Cl8 | C11H16Cl8 | 429.8731 | 428.8658 | 488.8869 | 16 |  | 11 |  |  | 7 | 1 |  |  |  |  | 11.7
| C11Cl9 | C11H15Cl9 | 463.8341 | 462.8268 | 522.848 | 15 |  | 11 |  |  | 8 | 1 |  |  |  |  | 12.3
| C11Cl10 | C11H14Cl10 | 497.7951 | 496.7879 | 556.809 | 14 |  | 11 |  |  | 9 | 1 |  |  |  |  | 12.9
| C11Cl11 | C11H13Cl11 | 531.7562 | 530.7489 | 590.77 | 13 |  | 11 |  |  | 10 | 1
| C12Cl4 | C12H22Cl4 | 308.0446 | 307.0373 | 367.0585 | 22 |  | 12 |  |  | 3 | 1 |  |  |  |  | 10.4
| C12Cl5 | C12H21Cl5 | 342.0056 | 340.9984 | 401.0195 | 21 |  | 12 |  |  | 4 | 1 |  |  |  |  | 10.7
| C12Cl6 | C12H20Cl6 | 375.9667 | 374.9594 | 434.9805 | 20 |  | 12 |  |  | 5 | 1 |  |  |  |  | 11.3
| C12Cl7 | C12H19Cl7 | 409.9277 | 408.9204 | 468.9415 | 19 |  | 12 |  |  | 6 | 1 |  |  |  |  | 11.8
| C12Cl8 | C12H18Cl8 | 443.8887 | 442.8814 | 502.9026 | 18 |  | 12 |  |  | 7 | 1 |  |  |  |  | 12.4
| C12Cl9 | C12H17Cl9 | 477.8497 | 476.8425 | 536.8636 | 17 |  | 12 |  |  | 8 | 1 |  |  |  |  | 12.9
| C12Cl10 | C12H16Cl10 | 511.8108 | 510.8035 | 570.8246 | 16 |  | 12 |  |  | 9 | 1 |  |  |  |  | 13.4
| C12Cl11 | C12H15Cl11 | 545.7718 | 544.7645 | 604.7857 | 15 |  | 12 |  |  | 10 | 1 |  |  |  |  | 13.8
| C12Cl12 | C12H14Cl12 | 579.7328 | 578.7256 | 638.7467 | 14 |  | 12 |  |  | 11 | 1
| C13Cl4 | C13H24Cl4 | 322.0603 | 321.053 | 381.0741 | 24 |  | 13 |  |  | 3 | 1 |  |  |  |  | 11.1
| C13Cl5 | C13H23Cl5 | 356.0213 | 355.014 | 415.0351 | 23 |  | 13 |  |  | 4 | 1 |  |  |  |  | 11.6
| C13Cl6 | C13H22Cl6 | 389.9823 | 388.975 | 448.9962 | 22 |  | 13 |  |  | 5 | 1 |  |  |  |  | 12.01
| C13Cl7 | C13H21Cl7 | 423.9433 | 422.9361 | 482.9572 | 21 |  | 13 |  |  | 6 | 1 |  |  |  |  | 12.5
| C13Cl8 | C13H20Cl8 | 457.9044 | 456.8971 | 516.9182 | 20 |  | 13 |  |  | 7 | 1 |  |  |  |  | 13
| C13Cl9 | C13H19Cl9 | 491.8654 | 490.8581 | 550.8793 | 19 |  | 13 |  |  | 8 | 1 |  |  |  |  | 13.5
| C13Cl10 | C13H18Cl10 | 525.8264 | 524.8192 | 584.8403 | 18 |  | 13 |  |  | 9 | 1 |  |  |  |  | 14.1
| C13Cl11 | C13H17Cl11 | 559.7875 | 558.7802 | 618.8013 | 17 |  | 13 |  |  | 10 | 1 |  |  |  |  | 14.4
| C13Cl12 | C13H16Cl12 | 593.7485 | 592.7412 | 652.7623 | 16 |  | 13 |  |  | 11 | 1
| C13Cl13 | C13H15Cl13 | 627.7095 | 626.7022 | 686.7234 | 15 |  | 13 |  |  | 12 | 1
MCCP | C14Cl4 | C14H26Cl4 | 336.0759 | 335.0686 | 395.0898 | 26 |  | 14 |  |  | 3 | 1 |  |  |  |  | 11.9
| C14Cl5 | C14H25Cl5 | 370.0369 | 369.0297 | 429.0508 | 25 |  | 14 |  |  | 4 | 1 |  |  |  |  | 12.3
| C14Cl6 | C14H24Cl6 | 403.998 | 402.9907 | 463.0118 | 24 |  | 14 |  |  | 5 | 1 |  |  |  |  | 12.7
| C14Cl7 | C14H23Cl7 | 437.959 | 436.9517 | 496.9728 | 23 |  | 14 |  |  | 6 | 1 |  |  |  |  | 13.1
| C14Cl8 | C14H22Cl8 | 471.92 | 470.9127 | 530.9339 | 22 |  | 14 |  |  | 7 | 1 |  |  |  |  | 13.5
| C14Cl9 | C14H21Cl9 | 505.881 | 504.8738 | 564.8949 | 21 |  | 14 |  |  | 8 | 1 |  |  |  |  | 14
| C14Cl10 | C14H20Cl10 | 539.8421 | 538.8348 | 598.8559 | 20 |  | 14 |  |  | 9 | 1 |  |  |  |  | 14.4
| C14Cl11 | C14H19Cl11 | 573.8031 | 572.7958 | 632.817 | 19 |  | 14 |  |  | 10 | 1 |  |  |  |  | 15
| C14Cl12 | C14H18Cl12 | 607.7641 | 606.7569 | 666.778 | 18 |  | 14 |  |  | 11 | 1
| C14Cl13 | C14H17Cl13 | 641.7252 | 640.7179 | 700.739 | 17 |  | 14 |  |  | 12 | 1
| C14Cl14 | C14H16Cl14 | 675.6862 | 674.6789 | 734.7 | 16 |  | 14 |  |  | 13 | 1
| C15Cl4 | C15H28Cl4 | 350.0916 | 349.0843 | 409.1054 | 28 |  | 15 |  |  | 3 | 1 |  |  |  |  | 12.6
| C15Cl5 | C15H27Cl5 | 384.0526 | 383.0453 | 443.0664 | 27 |  | 15 |  |  | 4 | 1 |  |  |  |  | 12.9
| C15Cl6 | C15H26Cl6 | 418.0136 | 417.0063 | 477.0275 | 26 |  | 15 |  |  | 5 | 1 |  |  |  |  | 13.3
| C15Cl7 | C15H25Cl7 | 451.9746 | 450.9674 | 510.9885 | 25 |  | 15 |  |  | 6 | 1 |  |  |  |  | 13.719
| C15Cl8 | C15H24Cl8 | 485.9357 | 484.9284 | 544.9495 | 24 |  | 15 |  |  | 7 | 1 |  |  |  |  | 14.2
| C15Cl9 | C15H23Cl9 | 519.8967 | 518.8894 | 578.9106 | 23 |  | 15 |  |  | 8 | 1 |  |  |  |  | 14.6
| C15Cl10 | C15H22Cl10 | 553.8577 | 552.8505 | 612.8716 | 22 |  | 15 |  |  | 9 | 1 |  |  |  |  | 15.1
| C15Cl11 | C15H21Cl11 | 587.8188 | 586.8115 | 646.8326 | 21 |  | 15 |  |  | 10 | 1 |  |  |  |  | 15.8
| C15Cl12 | C15H20Cl12 | 621.7798 | 620.7725 | 680.7936 | 20 |  | 15 |  |  | 11 | 1
| C15Cl13 | C15H19Cl13 | 655.7408 | 654.7335 | 714.7547 | 19 |  | 15 |  |  | 12 | 1
| C15Cl14 | C15H18Cl14 | 689.7018 | 688.6946 | 748.7157 | 18 |  | 15 |  |  | 13 | 1
| C15Cl15 | C15H17Cl15 | 723.6629 | 722.6556 | 782.6767 | 17 |  | 15 |  |  | 14 | 1
| C16Cl4 | C16H30Cl4 | 364.1072 | 363.0999 | 423.1211 | 30 |  | 16 |  |  | 3 | 1
| C16Cl5 | C16H29Cl5 | 398.0682 | 397.061 | 457.0821 | 29 |  | 16 |  |  | 4 | 1
| C16Cl6 | C16H28Cl6 | 432.0293 | 431.022 | 491.0431 | 28 |  | 16 |  |  | 5 | 1
| C16Cl7 | C16H27Cl7 | 465.9903 | 464.983 | 525.0041 | 27 |  | 16 |  |  | 6 | 1
| C16Cl8 | C16H26Cl8 | 499.9513 | 498.944 | 558.9652 | 26 |  | 16 |  |  | 7 | 1
| C16Cl9 | C16H25Cl9 | 533.9123 | 532.9051 | 592.9262 | 25 |  | 16 |  |  | 8 | 1
| C16Cl10 | C16H24Cl10 | 567.8734 | 566.8661 | 626.8872 | 24 |  | 16 |  |  | 9 | 1
| C16Cl11 | C16H23Cl11 | 601.8344 | 600.8271 | 660.8483 | 23 |  | 16 |  |  | 10 | 1
| C16Cl12 | C16H22Cl12 | 635.7954 | 634.7882 | 694.8093 | 22 |  | 16 |  |  | 11 | 1
| C16Cl13 | C16H21Cl13 | 669.7565 | 668.7492 | 728.7703 | 21 |  | 16 |  |  | 12 | 1
| C16Cl14 | C16H20Cl14 | 703.7175 | 702.7102 | 762.7313 | 20 |  | 16 |  |  | 13 | 1
| C16Cl15 | C16H19Cl15 | 737.6785 | 736.6712 | 796.6924 | 19 |  | 16 |  |  | 14 | 1
| C16Cl16 | C16H18Cl16 | 771.6395 | 770.6323 | 830.6534 | 18 |  | 16 |  |  | 15 | 1
| C17Cl4 | C17H32Cl4 | 378.1229 | 377.1156 | 437.1367 | 32 |  | 17 |  |  | 3 | 1
| C17Cl5 | C17H31Cl5 | 412.0839 | 411.0766 | 471.0977 | 31 |  | 17 |  |  | 4 | 1
| C17Cl6 | C17H30Cl6 | 446.0449 | 445.0376 | 505.0588 | 30 |  | 17 |  |  | 5 | 1
| C17Cl7 | C17H29Cl7 | 480.0059 | 478.9987 | 539.0198 | 29 |  | 17 |  |  | 6 | 1
| C17Cl8 | C17H28Cl8 | 513.967 | 512.9597 | 572.9808 | 28 |  | 17 |  |  | 7 | 1
| C17Cl9 | C17H27Cl9 | 547.928 | 546.9207 | 606.9419 | 27 |  | 17 |  |  | 8 | 1
| C17Cl10 | C17H26Cl10 | 581.889 | 580.8818 | 640.9029 | 26 |  | 17 |  |  | 9 | 1
| C17Cl11 | C17H25Cl11 | 615.8501 | 614.8428 | 674.8639 | 25 |  | 17 |  |  | 10 | 1
| C17Cl12 | C17H24Cl12 | 649.8111 | 648.8038 | 708.8249 | 24 |  | 17 |  |  | 11 | 1
| C17Cl13 | C17H23Cl13 | 683.7721 | 682.7648 | 742.786 | 23 |  | 17 |  |  | 12 | 1
| C17Cl14 | C17H22Cl14 | 717.7331 | 716.7259 | 776.747 | 22 |  | 17 |  |  | 13 | 1
| C17Cl15 | C17H21Cl15 | 751.6942 | 750.6869 | 810.708 | 21 |  | 17 |  |  | 14 | 1
| C17Cl16 | C17H20Cl16 | 785.6552 | 784.6479 | 844.669 | 20 |  | 17 |  |  | 15 | 1
| C17Cl17 | C17H19Cl17 | 819.6162 | 818.6089 | 878.6301 | 19 |  | 17 |  |  | 16 | 1
LCCP | C18Cl4 | C18H34Cl4 | 392.1385 | 391.1312 | 451.1524 | 34 |  | 18 |  |  | 3 | 1
| C18Cl5 | C18H33Cl5 | 426.0995 | 425.0923 | 485.1134 | 33 |  | 18 |  |  | 4 | 1
| C18Cl6 | C18H32Cl6 | 460.0606 | 459.0533 | 519.0744 | 32 |  | 18 |  |  | 5 | 1
| C18Cl7 | C18H31Cl7 | 494.0216 | 493.0143 | 553.0354 | 31 |  | 18 |  |  | 6 | 1
| C18Cl8 | C18H30Cl8 | 527.9826 | 526.9753 | 586.9965 | 30 |  | 18 |  |  | 7 | 1
| C18Cl9 | C18H29Cl9 | 561.9436 | 560.9364 | 620.9575 | 29 |  | 18 |  |  | 8 | 1
| C18Cl10 | C18H28Cl10 | 595.9047 | 594.8974 | 654.9185 | 28 |  | 18 |  |  | 9 | 1
| C18Cl11 | C18H27Cl11 | 629.8657 | 628.8584 | 688.8796 | 27 |  | 18 |  |  | 10 | 1
| C18Cl12 | C18H26Cl12 | 663.8267 | 662.8195 | 722.8406 | 26 |  | 18 |  |  | 11 | 1
| C18Cl13 | C18H25Cl13 | 697.7878 | 696.7805 | 756.8016 | 25 |  | 18 |  |  | 12 | 1
| C18Cl14 | C18H24Cl14 | 731.7488 | 730.7415 | 790.7626 | 24 |  | 18 |  |  | 13 | 1
| C18Cl15 | C18H23Cl15 | 765.7098 | 764.7025 | 824.7237 | 23 |  | 18 |  |  | 14 | 1
| C18Cl16 | C18H22Cl16 | 799.6708 | 798.6636 | 858.6847 | 22 |  | 18 |  |  | 15 | 1
| C18Cl17 | C18H21Cl17 | 833.6319 | 832.6246 | 892.6457 | 21 |  | 18 |  |  | 16 | 1
| C18Cl18 | C18H20Cl18 | 867.5929 | 866.5856 | 926.6068 | 20 |  | 18 |  |  | 17 | 1
| C19Cl4 | C19H36Cl4 | 406.1542 | 405.1469 | 465.168 | 36 |  | 19 |  |  | 3 | 1
| C19Cl5 | C19H35Cl5 | 440.1152 | 439.1079 | 499.129 | 35 |  | 19 |  |  | 4 | 1
| C19Cl6 | C19H34Cl6 | 474.0762 | 473.0689 | 533.0901 | 34 |  | 19 |  |  | 5 | 1
| C19Cl7 | C19H33Cl7 | 508.0372 | 507.03 | 567.0511 | 33 |  | 19 |  |  | 6 | 1
| C19Cl8 | C19H32Cl8 | 541.9983 | 540.991 | 601.0121 | 32 |  | 19 |  |  | 7 | 1
| C19Cl9 | C19H31Cl9 | 575.9593 | 574.952 | 634.9732 | 31 |  | 19 |  |  | 8 | 1
| C19Cl10 | C19H30Cl10 | 609.9203 | 608.9131 | 668.9342 | 30 |  | 19 |  |  | 9 | 1
| C19Cl11 | C19H29Cl11 | 643.8814 | 642.8741 | 702.8952 | 29 |  | 19 |  |  | 10 | 1
| C19Cl12 | C19H28Cl12 | 677.8424 | 676.8351 | 736.8562 | 28 |  | 19 |  |  | 11 | 1
| C19Cl13 | C19H27Cl13 | 711.8034 | 710.7961 | 770.8173 | 27 |  | 19 |  |  | 12 | 1
| C19Cl14 | C19H26Cl14 | 745.7644 | 744.7572 | 804.7783 | 26 |  | 19 |  |  | 13 | 1
| C19Cl15 | C19H25Cl15 | 779.7255 | 778.7182 | 838.7393 | 25 |  | 19 |  |  | 14 | 1
| C19Cl16 | C19H24Cl16 | 813.6865 | 812.6792 | 872.7003 | 24 |  | 19 |  |  | 15 | 1
| C19Cl17 | C19H23Cl17 | 847.6475 | 846.6402 | 906.6614 | 23 |  | 19 |  |  | 16 | 1
| C19Cl18 | C19H22Cl18 | 881.6085 | 880.6013 | 940.6224 | 22 |  | 19 |  |  | 17 | 1
| C19Cl19 | C19H21Cl19 | 915.5696 | 914.5623 | 974.5834 | 21 |  | 19 |  |  | 18 | 1
| C20Cl4 | C20H38Cl4 | 420.1698 | 419.1625 | 479.1837 | 38 |  | 20 |  |  | 3 | 1
| C20Cl5 | C20H37Cl5 | 454.1308 | 453.1236 | 513.1447 | 37 |  | 20 |  |  | 4 | 1
| C20Cl6 | C20H36Cl6 | 488.0919 | 487.0846 | 547.1057 | 36 |  | 20 |  |  | 5 | 1
| C20Cl7 | C20H35Cl7 | 522.0529 | 521.0456 | 581.0667 | 35 |  | 20 |  |  | 6 | 1
| C20Cl8 | C20H34Cl8 | 556.0139 | 555.0066 | 615.0278 | 34 |  | 20 |  |  | 7 | 1
| C20Cl9 | C20H33Cl9 | 589.975 | 588.9677 | 648.9888 | 33 |  | 20 |  |  | 8 | 1
| C20Cl10 | C20H32Cl10 | 623.936 | 622.9287 | 682.9498 | 32 |  | 20 |  |  | 9 | 1
| C20Cl11 | C20H31Cl11 | 657.897 | 656.8897 | 716.9109 | 31 |  | 20 |  |  | 10 | 1
| C20Cl12 | C20H30Cl12 | 691.858 | 690.8508 | 750.8719 | 30 |  | 20 |  |  | 11 | 1
| C20Cl13 | C20H29Cl13 | 725.8191 | 724.8118 | 784.8329 | 29 |  | 20 |  |  | 12 | 1
| C20Cl14 | C20H28Cl14 | 759.7801 | 758.7728 | 818.7939 | 28 |  | 20 |  |  | 13 | 1
| C20Cl15 | C20H27Cl15 | 793.7411 | 792.7338 | 852.755 | 27 |  | 20 |  |  | 14 | 1
| C20Cl16 | C20H26Cl16 | 827.7021 | 826.6949 | 886.716 | 26 |  | 20 |  |  | 15 | 1
| C20Cl17 | C20H25Cl17 | 861.6632 | 860.6559 | 920.677 | 25 |  | 20 |  |  | 16 | 1
| C20Cl18 | C20H24Cl18 | 895.6242 | 894.6169 | 954.6381 | 24 |  | 20 |  |  | 17 | 1
| C20Cl19 | C20H23Cl19 | 929.5852 | 928.578 | 988.5991 | 23 |  | 20 |  |  | 18 | 1
| C20Cl20 | C20H22Cl20 | 963.5463 | 962.539 | 1022.5601 | 22 |  | 20 |  |  | 19 | 1
| C21Cl4 | C21H40Cl4 | 434.1855 | 433.1782 | 493.1993 | 40 |  | 21 |  |  | 3 | 1
| C21Cl5 | C21H39Cl5 | 468.1465 | 467.1392 | 527.1603 | 39 |  | 21 |  |  | 4 | 1
| C21Cl6 | C21H38Cl6 | 502.1075 | 501.1002 | 561.1214 | 38 |  | 21 |  |  | 5 | 1
| C21Cl7 | C21H37Cl7 | 536.0685 | 535.0613 | 595.0824 | 37 |  | 21 |  |  | 6 | 1
| C21Cl8 | C21H36Cl8 | 570.0296 | 569.0223 | 629.0434 | 36 |  | 21 |  |  | 7 | 1
| C21Cl9 | C21H35Cl9 | 603.9906 | 602.9833 | 663.0045 | 35 |  | 21 |  |  | 8 | 1
| C21Cl10 | C21H34Cl10 | 637.9516 | 636.9444 | 696.9655 | 34 |  | 21 |  |  | 9 | 1
| C21Cl11 | C21H33Cl11 | 671.9127 | 670.9054 | 730.9265 | 33 |  | 21 |  |  | 10 | 1
| C21Cl12 | C21H32Cl12 | 705.8737 | 704.8664 | 764.8875 | 32 |  | 21 |  |  | 11 | 1
| C21Cl13 | C21H31Cl13 | 739.8347 | 738.8274 | 798.8486 | 31 |  | 21 |  |  | 12 | 1
| C21Cl14 | C21H30Cl14 | 773.7957 | 772.7885 | 832.8096 | 30 |  | 21 |  |  | 13 | 1
| C21Cl15 | C21H29Cl15 | 807.7568 | 806.7495 | 866.7706 | 29 |  | 21 |  |  | 14 | 1
| C21Cl16 | C21H28Cl16 | 841.7178 | 840.7105 | 900.7316 | 28 |  | 21 |  |  | 15 | 1
| C21Cl17 | C21H27Cl17 | 875.6788 | 874.6715 | 934.6927 | 27 |  | 21 |  |  | 16 | 1
| C21Cl18 | C21H26Cl18 | 909.6398 | 908.6326 | 968.6537 | 26 |  | 21 |  |  | 17 | 1
| C21Cl19 | C21H25Cl19 | 943.6009 | 942.5936 | 1002.6147 | 25 |  | 21 |  |  | 18 | 1
| C21Cl20 | C21H24Cl20 | 977.5619 | 976.5546 | 1036.5758 | 24 |  | 21 |  |  | 19 | 1
| C21Cl21 | C21H23Cl21 | 1011.5229 | 1010.5157 | 1070.5368 | 23 |  | 21 |  |  | 20 | 1
| C22Cl4 | C22H42Cl4 | 448.2011 | 447.1938 | 507.215 | 42 |  | 22 |  |  | 3 | 1
| C22Cl5 | C22H41Cl5 | 482.1621 | 481.1549 | 541.176 | 41 |  | 22 |  |  | 4 | 1
| C22Cl6 | C22H40Cl6 | 516.1232 | 515.1159 | 575.137 | 40 |  | 22 |  |  | 5 | 1
| C22Cl7 | C22H39Cl7 | 550.0842 | 549.0769 | 609.098 | 39 |  | 22 |  |  | 6 | 1
| C22Cl8 | C22H38Cl8 | 584.0452 | 583.0379 | 643.0591 | 38 |  | 22 |  |  | 7 | 1
| C22Cl9 | C22H37Cl9 | 618.0063 | 616.999 | 677.0201 | 37 |  | 22 |  |  | 8 | 1
| C22Cl10 | C22H36Cl10 | 651.9673 | 650.96 | 710.9811 | 36 |  | 22 |  |  | 9 | 1
| C22Cl11 | C22H35Cl11 | 685.9283 | 684.921 | 744.9422 | 35 |  | 22 |  |  | 10 | 1
| C22Cl12 | C22H34Cl12 | 719.8893 | 718.8821 | 778.9032 | 34 |  | 22 |  |  | 11 | 1
| C22Cl13 | C22H33Cl13 | 753.8504 | 752.8431 | 812.8642 | 33 |  | 22 |  |  | 12 | 1
| C22Cl14 | C22H32Cl14 | 787.8114 | 786.8041 | 846.8252 | 32 |  | 22 |  |  | 13 | 1
| C22Cl15 | C22H31Cl15 | 821.7724 | 820.7651 | 880.7863 | 31 |  | 22 |  |  | 14 | 1
| C22Cl16 | C22H30Cl16 | 855.7334 | 854.7262 | 914.7473 | 30 |  | 22 |  |  | 15 | 1
| C22Cl17 | C22H29Cl17 | 889.6945 | 888.6872 | 948.7083 | 29 |  | 22 |  |  | 16 | 1
| C22Cl18 | C22H28Cl18 | 923.6555 | 922.6482 | 982.6694 | 28 |  | 22 |  |  | 17 | 1
| C22Cl19 | C22H27Cl19 | 957.6165 | 956.6093 | 1016.6304 | 27 |  | 22 |  |  | 18 | 1
| C22Cl20 | C22H26Cl20 | 991.5776 | 990.5703 | 1050.5914 | 26 |  | 22 |  |  | 19 | 1
| C22Cl21 | C22H25Cl21 | 1025.5386 | 1024.5313 | 1084.5524 | 25 |  | 22 |  |  | 20 | 1
| C22Cl22 | C22H24Cl22 | 1059.4996 | 1058.4923 | 1118.5135 | 24 |  | 22 |  |  | 21 | 1
| C23Cl4 | C23H44Cl4 | 462.2168 | 461.2095 | 521.2306 | 44 |  | 23 |  |  | 3 | 1
| C23Cl5 | C23H43Cl5 | 496.1778 | 495.1705 | 555.1916 | 43 |  | 23 |  |  | 4 | 1
| C23Cl6 | C23H42Cl6 | 530.1388 | 529.1315 | 589.1527 | 42 |  | 23 |  |  | 5 | 1
| C23Cl7 | C23H41Cl7 | 564.0998 | 563.0926 | 623.1137 | 41 |  | 23 |  |  | 6 | 1
| C23Cl8 | C23H40Cl8 | 598.0609 | 597.0536 | 657.0747 | 40 |  | 23 |  |  | 7 | 1
| C23Cl9 | C23H39Cl9 | 632.0219 | 631.0146 | 691.0358 | 39 |  | 23 |  |  | 8 | 1
| C23Cl10 | C23H38Cl10 | 665.9829 | 664.9757 | 724.9968 | 38 |  | 23 |  |  | 9 | 1
| C23Cl11 | C23H37Cl11 | 699.944 | 698.9367 | 758.9578 | 37 |  | 23 |  |  | 10 | 1
| C23Cl12 | C23H36Cl12 | 733.905 | 732.8977 | 792.9188 | 36 |  | 23 |  |  | 11 | 1
| C23Cl13 | C23H35Cl13 | 767.866 | 766.8587 | 826.8799 | 35 |  | 23 |  |  | 12 | 1
| C23Cl14 | C23H34Cl14 | 801.827 | 800.8198 | 860.8409 | 34 |  | 23 |  |  | 13 | 1
| C23Cl15 | C23H33Cl15 | 835.7881 | 834.7808 | 894.8019 | 33 |  | 23 |  |  | 14 | 1
| C23Cl16 | C23H32Cl16 | 869.7491 | 868.7418 | 928.7629 | 32 |  | 23 |  |  | 15 | 1
| C23Cl17 | C23H31Cl17 | 903.7101 | 902.7028 | 962.724 | 31 |  | 23 |  |  | 16 | 1
| C23Cl18 | C23H30Cl18 | 937.6711 | 936.6639 | 996.685 | 30 |  | 23 |  |  | 17 | 1
| C23Cl19 | C23H29Cl19 | 971.6322 | 970.6249 | 1030.646 | 29 |  | 23 |  |  | 18 | 1
| C23Cl20 | C23H28Cl20 | 1005.5932 | 1004.5859 | 1064.6071 | 28 |  | 23 |  |  | 19 | 1
| C23Cl21 | C23H27Cl21 | 1039.5542 | 1038.547 | 1098.5681 | 27 |  | 23 |  |  | 20 | 1
| C23Cl22 | C23H26Cl22 | 1073.5153 | 1072.508 | 1132.5291 | 26 |  | 23 |  |  | 21 | 1
| C23Cl23 | C23H25Cl23 | 1107.4763 | 1106.469 | 1166.4901 | 25 |  | 23 |  |  | 22 | 1
| C24Cl4 | C24H46Cl4 | 476.2324 | 475.2251 | 535.2463 | 46 |  | 24 |  |  | 3 | 1
| C24Cl5 | C24H45Cl5 | 510.1934 | 509.1862 | 569.2073 | 45 |  | 24 |  |  | 4 | 1
| C24Cl6 | C24H44Cl6 | 544.1545 | 543.1472 | 603.1683 | 44 |  | 24 |  |  | 5 | 1
| C24Cl7 | C24H43Cl7 | 578.1155 | 577.1082 | 637.1293 | 43 |  | 24 |  |  | 6 | 1
| C24Cl8 | C24H42Cl8 | 612.0765 | 611.0692 | 671.0904 | 42 |  | 24 |  |  | 7 | 1
| C24Cl9 | C24H41Cl9 | 646.0376 | 645.0303 | 705.0514 | 41 |  | 24 |  |  | 8 | 1
| C24Cl10 | C24H40Cl10 | 679.9986 | 678.9913 | 739.0124 | 40 |  | 24 |  |  | 9 | 1
| C24Cl11 | C24H39Cl11 | 713.9596 | 712.9523 | 772.9735 | 39 |  | 24 |  |  | 10 | 1
| C24Cl12 | C24H38Cl12 | 747.9206 | 746.9134 | 806.9345 | 38 |  | 24 |  |  | 11 | 1
| C24Cl13 | C24H37Cl13 | 781.8817 | 780.8744 | 840.8955 | 37 |  | 24 |  |  | 12 | 1
| C24Cl14 | C24H36Cl14 | 815.8427 | 814.8354 | 874.8565 | 36 |  | 24 |  |  | 13 | 1
| C24Cl15 | C24H35Cl15 | 849.8037 | 848.7964 | 908.8176 | 35 |  | 24 |  |  | 14 | 1
| C24Cl16 | C24H34Cl16 | 883.7647 | 882.7575 | 942.7786 | 34 |  | 24 |  |  | 15 | 1
| C24Cl17 | C24H33Cl17 | 917.7258 | 916.7185 | 976.7396 | 33 |  | 24 |  |  | 16 | 1
| C24Cl18 | C24H32Cl18 | 951.6868 | 950.6795 | 1010.7007 | 32 |  | 24 |  |  | 17 | 1
| C24Cl19 | C24H31Cl19 | 985.6478 | 984.6406 | 1044.6617 | 31 |  | 24 |  |  | 18 | 1
| C24Cl20 | C24H30Cl20 | 1019.6089 | 1018.6016 | 1078.6227 | 30 |  | 24 |  |  | 19 | 1
| C24Cl21 | C24H29Cl21 | 1053.5699 | 1052.5626 | 1112.5837 | 29 |  | 24 |  |  | 20 | 1
| C24Cl22 | C24H28Cl22 | 1087.5309 | 1086.5236 | 1146.5448 | 28 |  | 24 |  |  | 21 | 1
| C24Cl23 | C24H27Cl23 | 1121.4919 | 1120.4847 | 1180.5058 | 27 |  | 24 |  |  | 22 | 1
| C24Cl24 | C24H26Cl24 | 1155.453 | 1154.4457 | 1214.4668 | 26 |  | 24 |  |  | 23 | 1
| C25Cl4 | C25H48Cl4 | 490.2481 | 489.2408 | 549.2619 | 48 |  | 25 |  |  | 3 | 1
| C25Cl5 | C25H47Cl5 | 524.2091 | 523.2018 | 583.2229 | 47 |  | 25 |  |  | 4 | 1
| C25Cl6 | C25H46Cl6 | 558.1701 | 557.1628 | 617.184 | 46 |  | 25 |  |  | 5 | 1
| C25Cl7 | C25H45Cl7 | 592.1311 | 591.1239 | 651.145 | 45 |  | 25 |  |  | 6 | 1
| C25Cl8 | C25H44Cl8 | 626.0922 | 625.0849 | 685.106 | 44 |  | 25 |  |  | 7 | 1
| C25Cl9 | C25H43Cl9 | 660.0532 | 659.0459 | 719.0671 | 43 |  | 25 |  |  | 8 | 1
| C25Cl10 | C25H42Cl10 | 694.0142 | 693.007 | 753.0281 | 42 |  | 25 |  |  | 9 | 1
| C25Cl11 | C25H41Cl11 | 727.9753 | 726.968 | 786.9891 | 41 |  | 25 |  |  | 10 | 1
| C25Cl12 | C25H40Cl12 | 761.9363 | 760.929 | 820.9501 | 40 |  | 25 |  |  | 11 | 1
| C25Cl13 | C25H39Cl13 | 795.8973 | 794.89 | 854.9112 | 39 |  | 25 |  |  | 12 | 1
| C25Cl14 | C25H38Cl14 | 829.8583 | 828.8511 | 888.8722 | 38 |  | 25 |  |  | 13 | 1
| C25Cl15 | C25H37Cl15 | 863.8194 | 862.8121 | 922.8332 | 37 |  | 25 |  |  | 14 | 1
| C25Cl16 | C25H36Cl16 | 897.7804 | 896.7731 | 956.7942 | 36 |  | 25 |  |  | 15 | 1
| C25Cl17 | C25H35Cl17 | 931.7414 | 930.7341 | 990.7553 | 35 |  | 25 |  |  | 16 | 1
| C25Cl18 | C25H34Cl18 | 965.7024 | 964.6952 | 1024.7163 | 34 |  | 25 |  |  | 17 | 1
| C25Cl19 | C25H33Cl19 | 999.6635 | 998.6562 | 1058.6773 | 33 |  | 25 |  |  | 18 | 1
| C25Cl20 | C25H32Cl20 | 1033.6245 | 1032.6172 | 1092.6384 | 32 |  | 25 |  |  | 19 | 1
| C25Cl21 | C25H31Cl21 | 1067.5855 | 1066.5783 | 1126.5994 | 31 |  | 25 |  |  | 20 | 1
| C25Cl22 | C25H30Cl22 | 1101.5466 | 1100.5393 | 1160.5604 | 30 |  | 25 |  |  | 21 | 1
| C25Cl23 | C25H29Cl23 | 1135.5076 | 1134.5003 | 1194.5214 | 29 |  | 25 |  |  | 22 | 1
| C25Cl24 | C25H28Cl24 | 1169.4686 | 1168.4613 | 1228.4825 | 28 |  | 25 |  |  | 23 | 1
| C25Cl25 | C25H27Cl25 | 1203.4296 | 1202.4224 | 1262.4435 | 27 |  | 25 |  |  | 24 | 1
| C26Cl4 | C26H50Cl4 | 504.2637 | 503.2564 | 563.2776 | 50 |  | 26 |  |  | 3 | 1
| C26Cl5 | C26H49Cl5 | 538.2247 | 537.2175 | 597.2386 | 49 |  | 26 |  |  | 4 | 1
| C26Cl6 | C26H48Cl6 | 572.1858 | 571.1785 | 631.1996 | 48 |  | 26 |  |  | 5 | 1
| C26Cl7 | C26H47Cl7 | 606.1468 | 605.1395 | 665.1606 | 47 |  | 26 |  |  | 6 | 1
| C26Cl8 | C26H46Cl8 | 640.1078 | 639.1005 | 699.1217 | 46 |  | 26 |  |  | 7 | 1
| C26Cl9 | C26H45Cl9 | 674.0689 | 673.0616 | 733.0827 | 45 |  | 26 |  |  | 8 | 1
| C26Cl10 | C26H44Cl10 | 708.0299 | 707.0226 | 767.0437 | 44 |  | 26 |  |  | 9 | 1
| C26Cl11 | C26H43Cl11 | 741.9909 | 740.9836 | 801.0048 | 43 |  | 26 |  |  | 10 | 1
| C26Cl12 | C26H42Cl12 | 775.9519 | 774.9447 | 834.9658 | 42 |  | 26 |  |  | 11 | 1
| C26Cl13 | C26H41Cl13 | 809.913 | 808.9057 | 868.9268 | 41 |  | 26 |  |  | 12 | 1
| C26Cl14 | C26H40Cl14 | 843.874 | 842.8667 | 902.8878 | 40 |  | 26 |  |  | 13 | 1
| C26Cl15 | C26H39Cl15 | 877.835 | 876.8277 | 936.8489 | 39 |  | 26 |  |  | 14 | 1
| C26Cl16 | C26H38Cl16 | 911.796 | 910.7888 | 970.8099 | 38 |  | 26 |  |  | 15 | 1
| C26Cl17 | C26H37Cl17 | 945.7571 | 944.7498 | 1004.7709 | 37 |  | 26 |  |  | 16 | 1
| C26Cl18 | C26H36Cl18 | 979.7181 | 978.7108 | 1038.732 | 36 |  | 26 |  |  | 17 | 1
| C26Cl19 | C26H35Cl19 | 1013.6791 | 1012.6719 | 1072.693 | 35 |  | 26 |  |  | 18 | 1
| C26Cl20 | C26H34Cl20 | 1047.6402 | 1046.6329 | 1106.654 | 34 |  | 26 |  |  | 19 | 1
| C26Cl21 | C26H33Cl21 | 1081.6012 | 1080.5939 | 1140.615 | 33 |  | 26 |  |  | 20 | 1
| C26Cl22 | C26H32Cl22 | 1115.5622 | 1114.5549 | 1174.5761 | 32 |  | 26 |  |  | 21 | 1
| C26Cl23 | C26H31Cl23 | 1149.5232 | 1148.516 | 1208.5371 | 31 |  | 26 |  |  | 22 | 1
| C26Cl24 | C26H30Cl24 | 1183.4843 | 1182.477 | 1242.4981 | 30 |  | 26 |  |  | 23 | 1
| C26Cl25 | C26H29Cl25 | 1217.4453 | 1216.438 | 1276.4591 | 29 |  | 26 |  |  | 24 | 1
| C26Cl26 | C26H28Cl26 | 1251.4063 | 1250.399 | 1310.4202 | 28 |  | 26 |  |  | 25 | 1
| C27Cl4 | C27H52Cl4 | 518.2794 | 517.2721 | 577.2932 | 52 |  | 27 |  |  | 3 | 1
| C27Cl5 | C27H51Cl5 | 552.2404 | 551.2331 | 611.2542 | 51 |  | 27 |  |  | 4 | 1
| C27Cl6 | C27H50Cl6 | 586.2014 | 585.1941 | 645.2153 | 50 |  | 27 |  |  | 5 | 1
| C27Cl7 | C27H49Cl7 | 620.1624 | 619.1552 | 679.1763 | 49 |  | 27 |  |  | 6 | 1
| C27Cl8 | C27H48Cl8 | 654.1235 | 653.1162 | 713.1373 | 48 |  | 27 |  |  | 7 | 1
| C27Cl9 | C27H47Cl9 | 688.0845 | 687.0772 | 747.0984 | 47 |  | 27 |  |  | 8 | 1
| C27Cl10 | C27H46Cl10 | 722.0455 | 721.0383 | 781.0594 | 46 |  | 27 |  |  | 9 | 1
| C27Cl11 | C27H45Cl11 | 756.0066 | 754.9993 | 815.0204 | 45 |  | 27 |  |  | 10 | 1
| C27Cl12 | C27H44Cl12 | 789.9676 | 788.9603 | 848.9814 | 44 |  | 27 |  |  | 11 | 1
| C27Cl13 | C27H43Cl13 | 823.9286 | 822.9213 | 882.9425 | 43 |  | 27 |  |  | 12 | 1
| C27Cl14 | C27H42Cl14 | 857.8896 | 856.8824 | 916.9035 | 42 |  | 27 |  |  | 13 | 1
| C27Cl15 | C27H41Cl15 | 891.8507 | 890.8434 | 950.8645 | 41 |  | 27 |  |  | 14 | 1
| C27Cl16 | C27H40Cl16 | 925.8117 | 924.8044 | 984.8255 | 40 |  | 27 |  |  | 15 | 1
| C27Cl17 | C27H39Cl17 | 959.7727 | 958.7654 | 1018.7866 | 39 |  | 27 |  |  | 16 | 1
| C27Cl18 | C27H38Cl18 | 993.7337 | 992.7265 | 1052.7476 | 38 |  | 27 |  |  | 17 | 1
| C27Cl19 | C27H37Cl19 | 1027.6948 | 1026.6875 | 1086.7086 | 37 |  | 27 |  |  | 18 | 1
| C27Cl20 | C27H36Cl20 | 1061.6558 | 1060.6485 | 1120.6697 | 36 |  | 27 |  |  | 19 | 1
| C27Cl21 | C27H35Cl21 | 1095.6168 | 1094.6096 | 1154.6307 | 35 |  | 27 |  |  | 20 | 1
| C27Cl22 | C27H34Cl22 | 1129.5779 | 1128.5706 | 1188.5917 | 34 |  | 27 |  |  | 21 | 1
| C27Cl23 | C27H33Cl23 | 1163.5389 | 1162.5316 | 1222.5527 | 33 |  | 27 |  |  | 22 | 1
| C27Cl24 | C27H32Cl24 | 1197.4999 | 1196.4926 | 1256.5138 | 32 |  | 27 |  |  | 23 | 1
| C27Cl25 | C27H31Cl25 | 1231.4609 | 1230.4537 | 1290.4748 | 31 |  | 27 |  |  | 24 | 1
| C27Cl26 | C27H30Cl26 | 1265.422 | 1264.4147 | 1324.4358 | 30 |  | 27 |  |  | 25 | 1
| C27Cl27 | C27H29Cl27 | 1299.383 | 1298.3757 | 1358.3969 | 29 |  | 27 |  |  | 26 | 1
| C28Cl4 | C28H54Cl4 | 532.295 | 531.2877 | 591.3089 | 54 |  | 28 |  |  | 3 | 1
| C28Cl5 | C28H53Cl5 | 566.256 | 565.2488 | 625.2699 | 53 |  | 28 |  |  | 4 | 1
| C28Cl6 | C28H52Cl6 | 600.2171 | 599.2098 | 659.2309 | 52 |  | 28 |  |  | 5 | 1
| C28Cl7 | C28H51Cl7 | 634.1781 | 633.1708 | 693.1919 | 51 |  | 28 |  |  | 6 | 1
| C28Cl8 | C28H50Cl8 | 668.1391 | 667.1318 | 727.153 | 50 |  | 28 |  |  | 7 | 1
| C28Cl9 | C28H49Cl9 | 702.1002 | 701.0929 | 761.114 | 49 |  | 28 |  |  | 8 | 1
| C28Cl10 | C28H48Cl10 | 736.0612 | 735.0539 | 795.075 | 48 |  | 28 |  |  | 9 | 1
| C28Cl11 | C28H47Cl11 | 770.0222 | 769.0149 | 829.0361 | 47 |  | 28 |  |  | 10 | 1
| C28Cl12 | C28H46Cl12 | 803.9832 | 802.976 | 862.9971 | 46 |  | 28 |  |  | 11 | 1
| C28Cl13 | C28H45Cl13 | 837.9443 | 836.937 | 896.9581 | 45 |  | 28 |  |  | 12 | 1
| C28Cl14 | C28H44Cl14 | 871.9053 | 870.898 | 930.9191 | 44 |  | 28 |  |  | 13 | 1
| C28Cl15 | C28H43Cl15 | 905.8663 | 904.859 | 964.8802 | 43 |  | 28 |  |  | 14 | 1
| C28Cl16 | C28H42Cl16 | 939.8273 | 938.8201 | 998.8412 | 42 |  | 28 |  |  | 15 | 1
| C28Cl17 | C28H41Cl17 | 973.7884 | 972.7811 | 1032.8022 | 41 |  | 28 |  |  | 16 | 1
| C28Cl18 | C28H40Cl18 | 1007.7494 | 1006.7421 | 1066.7633 | 40 |  | 28 |  |  | 17 | 1
| C28Cl19 | C28H39Cl19 | 1041.7104 | 1040.7032 | 1100.7243 | 39 |  | 28 |  |  | 18 | 1
| C28Cl20 | C28H38Cl20 | 1075.6715 | 1074.6642 | 1134.6853 | 38 |  | 28 |  |  | 19 | 1
| C28Cl21 | C28H37Cl21 | 1109.6325 | 1108.6252 | 1168.6463 | 37 |  | 28 |  |  | 20 | 1
| C28Cl22 | C28H36Cl22 | 1143.5935 | 1142.5862 | 1202.6074 | 36 |  | 28 |  |  | 21 | 1
| C28Cl23 | C28H35Cl23 | 1177.5545 | 1176.5473 | 1236.5684 | 35 |  | 28 |  |  | 22 | 1
| C28Cl24 | C28H34Cl24 | 1211.5156 | 1210.5083 | 1270.5294 | 34 |  | 28 |  |  | 23 | 1
| C28Cl25 | C28H33Cl25 | 1245.4766 | 1244.4693 | 1304.4904 | 33 |  | 28 |  |  | 24 | 1
| C28Cl26 | C28H32Cl26 | 1279.4376 | 1278.4303 | 1338.4515 | 32 |  | 28 |  |  | 25 | 1
| C28Cl27 | C28H31Cl27 | 1313.3986 | 1312.3914 | 1372.4125 | 31 |  | 28 |  |  | 26 | 1
| C28Cl28 | C28H30Cl28 | 1347.3597 | 1346.3524 | 1406.3735 | 30 |  | 28 |  |  | 27 | 1
| C29Cl4 | C29H56Cl4 | 546.3107 | 545.3034 | 605.3245 | 56 |  | 29 |  |  | 3 | 1
| C29Cl5 | C29H55Cl5 | 580.2717 | 579.2644 | 639.2855 | 55 |  | 29 |  |  | 4 | 1
| C29Cl6 | C29H54Cl6 | 614.2327 | 613.2254 | 673.2466 | 54 |  | 29 |  |  | 5 | 1
| C29Cl7 | C29H53Cl7 | 648.1937 | 647.1865 | 707.2076 | 53 |  | 29 |  |  | 6 | 1
| C29Cl8 | C29H52Cl8 | 682.1548 | 681.1475 | 741.1686 | 52 |  | 29 |  |  | 7 | 1
| C29Cl9 | C29H51Cl9 | 716.1158 | 715.1085 | 775.1297 | 51 |  | 29 |  |  | 8 | 1
| C29Cl10 | C29H50Cl10 | 750.0768 | 749.0696 | 809.0907 | 50 |  | 29 |  |  | 9 | 1
| C29Cl11 | C29H49Cl11 | 784.0379 | 783.0306 | 843.0517 | 49 |  | 29 |  |  | 10 | 1
| C29Cl12 | C29H48Cl12 | 817.9989 | 816.9916 | 877.0127 | 48 |  | 29 |  |  | 11 | 1
| C29Cl13 | C29H47Cl13 | 851.9599 | 850.9526 | 910.9738 | 47 |  | 29 |  |  | 12 | 1
| C29Cl14 | C29H46Cl14 | 885.9209 | 884.9137 | 944.9348 | 46 |  | 29 |  |  | 13 | 1
| C29Cl15 | C29H45Cl15 | 919.882 | 918.8747 | 978.8958 | 45 |  | 29 |  |  | 14 | 1
| C29Cl16 | C29H44Cl16 | 953.843 | 952.8357 | 1012.8568 | 44 |  | 29 |  |  | 15 | 1
| C29Cl17 | C29H43Cl17 | 987.804 | 986.7967 | 1046.8179 | 43 |  | 29 |  |  | 16 | 1
| C29Cl18 | C29H42Cl18 | 1021.765 | 1020.7578 | 1080.7789 | 42 |  | 29 |  |  | 17 | 1
| C29Cl19 | C29H41Cl19 | 1055.7261 | 1054.7188 | 1114.7399 | 41 |  | 29 |  |  | 18 | 1
| C29Cl20 | C29H40Cl20 | 1089.6871 | 1088.6798 | 1148.701 | 40 |  | 29 |  |  | 19 | 1
| C29Cl21 | C29H39Cl21 | 1123.6481 | 1122.6409 | 1182.662 | 39 |  | 29 |  |  | 20 | 1
| C29Cl22 | C29H38Cl22 | 1157.6092 | 1156.6019 | 1216.623 | 38 |  | 29 |  |  | 21 | 1
| C29Cl23 | C29H37Cl23 | 1191.5702 | 1190.5629 | 1250.584 | 37 |  | 29 |  |  | 22 | 1
| C29Cl24 | C29H36Cl24 | 1225.5312 | 1224.5239 | 1284.5451 | 36 |  | 29 |  |  | 23 | 1
| C29Cl25 | C29H35Cl25 | 1259.4922 | 1258.485 | 1318.5061 | 35 |  | 29 |  |  | 24 | 1
| C29Cl26 | C29H34Cl26 | 1293.4533 | 1292.446 | 1352.4671 | 34 |  | 29 |  |  | 25 | 1
| C29Cl27 | C29H33Cl27 | 1327.4143 | 1326.407 | 1386.4282 | 33 |  | 29 |  |  | 26 | 1
| C29Cl28 | C29H32Cl28 | 1361.3753 | 1360.368 | 1420.3892 | 32 |  | 29 |  |  | 27 | 1
| C29Cl29 | C29H31Cl29 | 1395.3364 | 1394.3291 | 1454.3502 | 31 |  | 29 |  |  | 28 | 1
| C30Cl4 | C30H58Cl4 | 560.3263 | 559.319 | 619.3402 | 58 |  | 30 |  |  | 3 | 1
| C30Cl5 | C30H57Cl5 | 594.2873 | 593.2801 | 653.3012 | 57 |  | 30 |  |  | 4 | 1
| C30Cl6 | C30H56Cl6 | 628.2484 | 627.2411 | 687.2622 | 56 |  | 30 |  |  | 5 | 1
| C30Cl7 | C30H55Cl7 | 662.2094 | 661.2021 | 721.2232 | 55 |  | 30 |  |  | 6 | 1
| C30Cl8 | C30H54Cl8 | 696.1704 | 695.1631 | 755.1843 | 54 |  | 30 |  |  | 7 | 1
| C30Cl9 | C30H53Cl9 | 730.1315 | 729.1242 | 789.1453 | 53 |  | 30 |  |  | 8 | 1
| C30Cl10 | C30H52Cl10 | 764.0925 | 763.0852 | 823.1063 | 52 |  | 30 |  |  | 9 | 1
| C30Cl11 | C30H51Cl11 | 798.0535 | 797.0462 | 857.0674 | 51 |  | 30 |  |  | 10 | 1
| C30Cl12 | C30H50Cl12 | 832.0145 | 831.0073 | 891.0284 | 50 |  | 30 |  |  | 11 | 1
| C30Cl13 | C30H49Cl13 | 865.9756 | 864.9683 | 924.9894 | 49 |  | 30 |  |  | 12 | 1
| C30Cl14 | C30H48Cl14 | 899.9366 | 898.9293 | 958.9504 | 48 |  | 30 |  |  | 13 | 1
| C30Cl15 | C30H47Cl15 | 933.8976 | 932.8903 | 992.9115 | 47 |  | 30 |  |  | 14 | 1
| C30Cl16 | C30H46Cl16 | 967.8586 | 966.8514 | 1026.8725 | 46 |  | 30 |  |  | 15 | 1
| C30Cl17 | C30H45Cl17 | 1001.8197 | 1000.8124 | 1060.8335 | 45 |  | 30 |  |  | 16 | 1
| C30Cl18 | C30H44Cl18 | 1035.7807 | 1034.7734 | 1094.7946 | 44 |  | 30 |  |  | 17 | 1
| C30Cl19 | C30H43Cl19 | 1069.7417 | 1068.7345 | 1128.7556 | 43 |  | 30 |  |  | 18 | 1
| C30Cl20 | C30H42Cl20 | 1103.7028 | 1102.6955 | 1162.7166 | 42 |  | 30 |  |  | 19 | 1
| C30Cl21 | C30H41Cl21 | 1137.6638 | 1136.6565 | 1196.6776 | 41 |  | 30 |  |  | 20 | 1
| C30Cl22 | C30H40Cl22 | 1171.6248 | 1170.6175 | 1230.6387 | 40 |  | 30 |  |  | 21 | 1
| C30Cl23 | C30H39Cl23 | 1205.5858 | 1204.5786 | 1264.5997 | 39 |  | 30 |  |  | 22 | 1
| C30Cl24 | C30H38Cl24 | 1239.5469 | 1238.5396 | 1298.5607 | 38 |  | 30 |  |  | 23 | 1
| C30Cl25 | C30H37Cl25 | 1273.5079 | 1272.5006 | 1332.5217 | 37 |  | 30 |  |  | 24 | 1
| C30Cl26 | C30H36Cl26 | 1307.4689 | 1306.4616 | 1366.4828 | 36 |  | 30 |  |  | 25 | 1
| C30Cl27 | C30H35Cl27 | 1341.4299 | 1340.4227 | 1400.4438 | 35 |  | 30 |  |  | 26 | 1
| C30Cl28 | C30H34Cl28 | 1375.391 | 1374.3837 | 1434.4048 | 34 |  | 30 |  |  | 27 | 1
| C30Cl29 | C30H33Cl29 | 1409.352 | 1408.3447 | 1468.3659 | 33 |  | 30 |  |  | 28 | 1
| C30Cl30 | C30H32Cl30 | 1443.313 | 1442.3058 | 1502.3269 | 32 |  | 30 |  |  | 29 | 1

## LC-TOF method
LC-TOF測定条件
UHPLC system | Agilent 1290 Infinity LC system
Column | ZORBAX RRHD SB-CN 2.1 x 100 mm, 1.8 um
Column Temp. | 40 deg C
Mobile phase A: | 5 mM CH3COONH4 aq
Mobile phase B: | 5 mM CH3COONH4 in methanol
Flow Rate | 0.4 mL/min
Injection volume | 5 uL
Gradient parameter
Time (min) | %B | Flow rate (mL/min)
0 | 45 | 0.4
12 | 75 | 0.4 | 2.5
22 | 90 | 0.4 | 1.5
32 | 99 | 0.4 | 0.9
35 | 99 | 0.4
35.1 | 45 | 0.4
40 | 45 | 0.4
QTOF system | Agilent 6530 True High Definition Q-TOF mass spectrometer
Ionization | Agilent Jet Stream ESI negative mode
Gas Temp. | 250 °C
Drying gas | 10 L/min
Nebulizer | 50 psig
Sheath gas Temp. | 250 °C
Sheath gas flow | 6 L/min
VCap | 3000 V
Nozzle voltage | 50 V
Fragmentor | 100 V
Skimmer | 65 V
OCT 1 RF Vpp | 750 V
Mass range | m/z 150 to 1600
Acquisition time | 1 spectra/s

## CP含有製品の試料前処理フロー
CP含有製品の試料前処理フロー
試料の撮影
ハンドヘルドXRF(HH-XRF)を用いた塩素含有量測定
試料の粗破砕 |  | 3ｇ
試料の凍結粉砕
試料 |  | 0.2 g
超音波抽出 |  | トルエン10 mL/30min
抽出液分取 |  | 0.5ml/10 mL
←10%ジクロロメタン/ヘキサン 3 mL
←55%硫酸シリカゲル 1 g
10%ジクロロメタン/ヘキサン |  |  | 55%硫酸シリカゲル
←10%ジクロロメタン/ヘキサン 3 mL
|  | 10%ジクロロメタン/ヘキサン |  |  | 55%硫酸シリカゲル
←10%ジクロロメタン/ヘキサン 3 mL
|  |  |  | 10%ジクロロメタン/ヘキサン |  |  | 55%硫酸シリカゲル
ろ過
濃縮 |  | 0.1 mL
定容 |  | メタノール 1 mL
フィルタろ過
IS添加 |  | D18-HBCD 1000 ng/mL x 0.05 mL
測定 |  | LC-ESI-MSMS

## 検量線溶液中の同族体組成
検量線溶液中の同族体組成
|  | Chloroparaffin C10-C13 51.5%Cl | Chloroparaffin C10-C13 63%Cl | SCCP-MIX-CS (1:1 mixture) |  |  | Chloroparaffin C14-C17 42%Cl | Chloroparaffin C14-C17 57%Cl | MCCP-MIX-CS (1:1 mixture) |  |  | Chloroparaffin C18-C20 36%Cl | Chloroparaffin C18-C20 49%Cl | LCCP-MIX-CS (1:1 mixture)
|  | % | % | % |  |  | % | % | % |  |  | % | % | %
C10 | Cl4 | 2.1 | 0.056 | 1.1 | C14 | Cl4 | 19.15 | 0.09 | 9.9 | C18 | Cl4 | 12.2962 | 0.28875 | 8.9 | 2.156 | 19.24 | 12.5849
| Cl5 | 2.6 | 0.82 | 1.7 |  | Cl5 | 28.41 | 2.11 | 16 |  | Cl5 | 19.1488 | 2.2584 | 15 | 3.42 | 30.52 | 21.4072
| Cl6 | 0.67 | 3.1 | 1.9 |  | Cl6 | 15.72 | 11.43 | 14 |  | Cl6 | 15.022 | 9.0519 | 17 | 3.77 | 27.15 | 24.0739
| Cl7 | 0.093 | 2.9 | 1.5 |  | Cl7 | 3.47 | 22.75 | 13 |  | Cl7 | 5.6821 | 15.565 | 15 | 2.993 | 26.22 | 21.2471
| Cl8 | 0 | 1 | 0.5 |  | Cl8 | 0.43 | 20 | 11 |  | Cl8 | 1.1362 | 12.8893 | 10 | 1 | 20.43 | 14.0255
| Cl9 | 0 | 0.17 | 0.085 |  | Cl9 | 0.02 | 7.8 | 4 |  | Cl9 | 0.17242 | 5.231 | 3.8 | 0.17 | 7.82 | 5.4034
| Cl10 | 0 | 0 | 0 |  | Cl10 | 0 | 1.34 | 0.69 |  | Cl10 | 0.021287 | 1.0547 | 0.76 | 0 | 1.34 | 1.076
C11 | Cl4 | 11 | 0.047 | 5.5 | C15 | Cl4 | 4.91 | 0.01 | 2.5 | C19 | Cl4 | 3.2702 | 0.053611 | 2.4 | 11.047 | 4.92 | 3.3238
| Cl5 | 15 | 1.3 | 8.2 |  | Cl5 | 9.53 | 0.34 | 5.1 |  | Cl5 | 5.8064 | 0.50278 | 4.5 | 16.3 | 9.87 | 6.3092
| Cl6 | 6.9 | 6.7 | 6.8 |  | Cl6 | 6.59 | 2.45 | 4.7 |  | Cl6 | 5.3543 | 2.33 | 5.5 | 13.6 | 9.04 | 7.6842
| Cl7 | 1.3 | 13 | 7.2 |  | Cl7 | 2.04 | 7.3 | 4.8 |  | Cl7 | 2.2673 | 4.9971 | 5.2 | 14.3 | 9.34 | 7.2645
| Cl8 | 0.13 | 9.9 | 5 |  | Cl8 | 0.26 | 9.03 | 4.8 |  | Cl8 | 0.50994 | 5.0578 | 4 | 10.03 | 9.29 | 5.5677
| Cl9 | 0 | 2.9 | 1.5 |  | Cl9 | 0.02 | 4.99 | 2.6 |  | Cl9 | 0.080203 | 2.4101 | 1.8 | 2.9 | 5.01 | 2.4904
| Cl10 | 0 | 0.27 | 0.14 |  | Cl10 | 0 | 1.21 | 0.62 |  | Cl10 | 0.0094457 | 0.58683 | 0.42 | 0.27 | 1.21 | 0.59627
C12 | Cl4 | 7.7 | 0 | 3.9 | C16 | Cl4 | 0.57 | 0 | 0.29 | C20 | Cl4 | 0.65803 | 0 | 0.47 | 7.7 | 0.57 | 0.65803
| Cl5 | 14 | 0.28 | 7.1 |  | Cl5 | 1.37 | 0 | 0.7 |  | Cl5 | 1.3553 | 0.052352 | 1 | 14.28 | 1.37 | 1.4076
| Cl6 | 11 | 2.8 | 6.9 |  | Cl6 | 1.16 | 0.34 | 0.77 |  | Cl6 | 1.4635 | 0.4001 | 1.3 | 13.8 | 1.5 | 1.8636
| Cl7 | 2.9 | 8.3 | 5.6 |  | Cl7 | 0.42 | 1.47 | 0.97 |  | Cl7 | 0.69921 | 1.0247 | 1.2 | 11.2 | 1.89 | 1.7239
| Cl8 | 0.46 | 12 | 6.2 |  | Cl8 | 0.06 | 2.51 | 1.3 |  | Cl8 | 0.18376 | 1.1892 | 0.97 | 12.46 | 2.57 | 1.373
| Cl9 | 0.029 | 7.1 | 3.6 |  | Cl9 | 0 | 1.88 | 0.97 |  | Cl9 | 0.039648 | 0.64671 | 0.49 | 7.129 | 1.88 | 0.68635
| Cl10 | 0 | 1.4 | 0.7 |  | Cl10 | 0 | 0.6 | 0.31 |  | Cl10 | 0 | 0.17677 | 0.13 | 1.4 | 0.6 | 0.17677
C13 | Cl4 | 2.8 | 0 | 1.4 | C17 | Cl4 | 0.06 | 0 | 0.031 | [5 empty] | 2.8 | 0.06
| Cl5 | 7.8 | 0 | 3.9 |  | Cl5 | 0.19 | 0 | 0.098 | [5 empty] | 7.8 | 0.19
| Cl6 | 9 | 0.46 | 4.7 |  | Cl6 | 0.19 | 0 | 0.098 | [5 empty] | 9.46 | 0.19
| Cl7 | 3.91 | 4.1 | 4 |  | Cl7 | 0.08 | 0.32 | 0.21 | [5 empty] | 8.01 | 0.4
| Cl8 | 0.88 | 9.6 | 5.2 |  | Cl8 | 0.01 | 0.74 | 0.39 | [5 empty] | 10.48 | 0.75
| Cl9 | 0.077 | 8.2 | 4.1 |  | Cl9 | 0 | 0.7 | 0.36 | [5 empty] | 8.277 | 0.7
| Cl10 | 0 | 3 | 1.5 |  | Cl10 | 0 | 0.26 | 0.13 | [5 empty] | 3 | 0.26

## 塩素同位体パターン
塩素同位体パターン
| M | M+1 | M+2 | M+3 | M+4 | M+5 | M+6 | M+7 | M+8 | M+9 | M+10 | M+11 | M+12 | M+13 | M+14
4Cl | 2943000000000 | 335300000000 | 3782000000000 | 429400000000 | 1828000000000 | 206400000000 | 395600000000
5Cl | 2231000000000 | 253400000000 | 3579000000000 | 405600000000 | 2302000000000 | 259800000000 | 742700000000 |  | 120900000000
6Cl | 1690000000000 | 191800000000 | 3253000000000 | 368300000000 | 2612000000000 | 294800000000 | 1121000000000 | 125900000000 | 271600000000
7Cl | 1281000000000 | 145200000000 | 2875000000000 | 325200000000 | 2767000000000 | 312200000000 | 1482000000000 | 166600000000 | 477400000000
8Cl | 970700000000 | 109900000000 | 2489000000000 | 281200000000 | 2794000000000 | 315000000000 | 1794000000000 | 201700000000 | 720900000000 |  | 185900000000
9Cl | 671700000000 | 137300000000 | 1947000000000 | 396000000000 | 2511000000000 | 507800000000 | 1894000000000 | 380100000000 | 921700000000 | 183100000000 | 300600000000
10Cl | 509000000000 | 104000000000 | 1638000000000 | 333100000000 | 2375000000000 | 480400000000 | 2044000000000 | 410800000000 | 1157000000000 | 230700000000 | 451100000000 |  | 122800000000
11Cl | 422400000000 |  | 1488000000000 | 167800000000 | 2384000000000 | 268400000000 | 2293000000000 | 257700000000 | 1471000000000 | 164900000000 | 661000000000 |  | 212400000000
12Cl | 316500000000 |  | 1217000000000 | 150800000000 | 2145000000000 | 265500000000 | 2292000000000 | 283200000000 | 1654000000000 | 203900000000 | 849700000000 | 104500000000 | 318600000000
13Cl | 237100000000 |  | 987800000000 | 133600000000 | 1900000000000 | 256400000000 | 2234000000000 | 300900000000 | 1792000000000 | 240800000000 | 1036000000000 | 138800000000 | 443800000000 |  | 142800000000
14Cl | 177700000000 |  | 797200000000 | 116800000000 | 1661000000000 | 242900000000 | 2131000000000 | 310900000000 | 1881000000000 | 273700000000 | 1208000000000 | 175300000000 | 582400000000 |  | 214200000000
| M | M+1 | M+2 | M+3 | M+4 | M+5 | M+6 | M+7 | M+8 | M+9 | M+10 | M+11 | M+12 | M+13 | M+14
4Cl | 0.29668 | 0.033801 | 0.38126 | 0.043288 | 0.18428 | 0.020807 | 0.03988
5Cl | 0.22548 | 0.02561 | 0.36172 | 0.040993 | 0.23266 | 0.026257 | 0.075063 |  | 0.012219
6Cl | 0.17022 | 0.019318 | 0.32765 | 0.037096 | 0.26308 | 0.029693 | 0.11291 | 0.012681 | 0.027356
7Cl | 0.13029 | 0.014769 | 0.29242 | 0.033077 | 0.28144 | 0.031755 | 0.15074 | 0.016945 | 0.048558
8Cl | 0.098425 | 0.011143 | 0.25238 | 0.028513 | 0.2833 | 0.03194 | 0.1819 | 0.020452 | 0.073097 |  | 0.01885
9Cl | 0.068191 | 0.013939 | 0.19766 | 0.040202 | 0.25492 | 0.051552 | 0.19228 | 0.038588 | 0.093571 | 0.018588 | 0.030517
10Cl | 0.051644 | 0.010552 | 0.16619 | 0.033797 | 0.24097 | 0.048742 | 0.20739 | 0.041681 | 0.11739 | 0.023407 | 0.04577 |  | 0.01246
11Cl | 0.043143 |  | 0.15198 | 0.017139 | 0.2435 | 0.027414 | 0.2342 | 0.026321 | 0.15025 | 0.016843 | 0.067514 |  | 0.021694
12Cl | 0.032294 |  | 0.12417 | 0.015387 | 0.21886 | 0.02709 | 0.23386 | 0.028896 | 0.16876 | 0.020805 | 0.086698 | 0.010663 | 0.032508
13Cl | 0.02444 |  | 0.10182 | 0.013771 | 0.19585 | 0.02643 | 0.23028 | 0.031017 | 0.18472 | 0.024822 | 0.10679 | 0.014308 | 0.045747 |  | 0.01472
14Cl | 0.018592 |  | 0.083407 | 0.01222 | 0.17378 | 0.025414 | 0.22296 | 0.032528 | 0.1968 | 0.028636 | 0.12639 | 0.018341 | 0.060934 |  | 0.022411